Porosimetric membrane characterization techniques: a review

Melike Begum Tanis-Kanbura,b,c, René I. Peinadord, José I. Calvoe, Antonio Hernándeze, Xiao Huf,g, Jia Wei Chewa,c,*

a School of Chemical and Biomedical Engineering, Nanyang Technological University, 637459, Singapore

b Interdisciplinary Graduate School, Nanyang Technological University, 639798, Singapore

c Singapore Membrane Technology Centre, Nanyang Environment and Water Research Institute, Nanyang Technological University, 637141, Singapore

d Institut de la Filtration et des Techniques Séparatives (IFTS), Rue Marcel Pagnol, 47510, Foulayronnes, France

e Surfaces and Porous Materials (SMAP), Associated Research Unit to CSIC, UVainnova bldg, Pº de Belén 11 and Institute of Sustainable Processes (ISP), Dr. Mergelina s/n, University of Valladolid, 47071, Valladolid, Spain.

f Environmental Chemistry and Materials Centre, Nanyang Environment and Water Research Institute, Nanyang Technological University, Singapore 637141, Singapore

g School of Materials Science and Engineering, Nanyang Technological University, Singapore 639798, Singapore

* Corresponding author. School of Chemical and Biomedical Engineering, Nanyang Technological University, 637459, Singapore.

E-mail address: JChew@ntu.edu.sg (J. W. Chew).

# Abstract

Membrane technology has been significantly gaining importance in water treatment applications recently due to their environmentally friendly operation, and less complex and cheaper operating conditions compared to the conventional water treatment studies. To provide sustainable and efficient membrane applications, membrane selection is a crucial factor for both fabrication and operation processes, and it is determined by membrane characterization, which presents critical parameters such as average membrane pore size and pore size distribution (PSD). There are two main characterization methods as the direct and indirect, the latter is based on mathematical equations, more affordable, and it can characterize larger membrane areas compared to the direct techniques. This study reviews the indirect membrane characterization methods according to their backgrounds that are Young-Laplace Equation-based, Kelvin Equation-based, Gibbs-Thomson Equation-based, and spectroscopy-based techniques. After giving the theoretical details of the methods, the measurement details and relevant experimental procedures are presented. Lastly, the remarkable indirect membrane studies, which have been performed from past to present, are explained in detail.

Keywords: Membrane characterization; pore-size distribution (PSD); porosity; water treatment; polymeric membranes; ceramic membranes.

Contents

Abstract	1

1.	Introduction	4

2.	Indirect Method Theories	9

2.1.	Young-Laplace Equation	9

2.2.	Kelvin Equation	10

2.3.	Gibbs-Thomson Equation	13

2.4.	Spectroscopy-Based Theories	16

2.4.1.	Positron Annihilation Lifetime Spectroscopy (PALS)	16

2.4.2.	Synchrotron Radiation (SR)	18

3.	Young-Laplace Equation-Based Techniques	18

3.1.	Liquid Displacement Porosimetry (LDP)	18

3.1.1.	Bubble-Point Test	18

3.1.2.	Gas-Liquid Displacement Porosimetry (GLDP)	20

3.1.3.	Liquid-Liquid Displacement Porosimetry (LLDP)	29

3.2.	Mercury Intrusion Porosimetry (HgP)	46

4.	Kelvin Equation-Based Techniques	51

4.1.	Evapoporometry (EP)	51

4.2.	Permporometry (PmP)	60

4.3.	Gas Adsorption-Desorption (GAD)	69

4.3.1.	Determination of Surface Area	70

4.3.2.	Determination of Mean Pore Size and Pore Size Distribution	71

4.3.3.	Determination of Overall Porosity	74

5.	Gibbs-Thomson Equation-Based Techniques	76

5.1.	Thermoporometry (ThP)	76

6.	Spectroscopy-Based Techniques	78

6.1.	Positron Annihilation Lifetime Spectroscopy (PALS)	79

6.2.	Synchrotron Radiation (SR)	81

7.	Conclusions and Future Directions	83

Acknowledgments	86

References	86

Introduction

Porous materials have widespread applications in many industrial and research fields. In particular, membranes are simple, easy-to-operate, and allow for environmentally friendly applications with inexpensive operating costs and fewer effluents; hence, they have been gaining importance as an affordable and feasible alternative to conventional separation means [1]. Loeb and Sourirajan [2] developed the first asymmetric membrane in the early 1960s, unveiling the complex procedures to overcome the challenges encountered during membrane fabrication. Apart from the cellulose acetate membrane that made seawater desalination through reverse osmosis the most dominant method, innumerable processes in which membranes are used for wide-ranging separation have sprouted.

To ensure selectivity and permeability in a wide range of separations, membranes may differ significantly in structure and constitution, and consequently in their functional behaviours. Therefore, it is necessary to adequately characterize the membranes so that they can be selected and/or adjusted for each existing or potential application. This knowledge is also important for membrane manufacturers to optimize the fabrication parameters during membrane production for the most suitable membrane for a given separation. Besides, membrane fouling, which causes decrements in permeate flux and lifetime over time, results in a significant disadvantage for the operations of the membranes. The analysis and selection of convenient membrane type are critical points for the membrane-based applications. Hence, characterization is a fundamental step in membrane production and one of the key factors to better understand the fouling behaviour and stability of the membrane [3]. But how to define this term? In a generalist way, we can define characterization of a membrane as “the most complete knowledge of its constitution, structure and functional behaviour, obtained through the combined and critical use of adequate methods and techniques”[4]. Accordingly, it is stated that the main aim being pursued in the characterization of a membrane is the prediction of its performance when used in a given process [4].

In most of the pioneering membrane applications, the size of the pores present in a membrane is a key factor in determining its possible applications and capabilities. The characterization of pore size includes the determination of (average) pore size diameter and the pore size distribution (PSD). The average pore size of the membrane filter can determine, in the first instance, the size of the molecules to be retained. But more complete knowledge would be acquired if we supply not only an average or mean pore size but also the complete PSD, let’s say, the amount or percentage for each pore size of those present in the membrane. Particle removal performance can be simply analyzed by using the PSD trends of a membrane. There are two main classification groups for membrane pore size characterization which are direct and indirect methods, respectively.

Several techniques can be considered as direct methods since they provide live images to observe the membrane morphological details such as cross-sectional area, surface porosity, pore size, pore shape, and PSD [5], and measure them by some sort of image analysis and spectrometric methods. This group can include all techniques based on microscopic inspection of surface or cross-sections of the membranes such as scanning electron microscopy (SEM), field emission scanning electron microscopy (FESEM), scanning tunneling microscopy (STM), atomic force microscopy (AFM), and transmission electron microscopy (TEM) [6-9]; and the spectroscopic ones such as electrochemical impedance spectroscopy (EIS), Raman spectroscopy (RS), and nuclear magnetic resonance (NMR) [10-23]. Solute retention test is a characterization method that using appropriate modeling to convert retention of selected molecules into PSD is more related to the knowledge about membrane performance than with actual filter structure. Although the direct methods provide benefits such as live observation, they also bring some drawbacks for the pore size characterization such as very small area observation and coating with a conductive material (especially for SEM and FESEM) [24]. To deal with the mentioned drawbacks, the indirect characterization methods, which are based on some theoretical basis to convert direct measurements into pore size distributions [9], are also preferred for the membrane characterization.

The indirect techniques give such complete information about PSD can be grouped under the term porometries (or porosimetries which perhaps sounds more concerned with the whole membrane porosity). The techniques developed in the past centuries (mostly during the 20th century) can give a more or less complete and accurate manner for the PSD determination of a given membrane filter. They can be distinguished in the base of the physical phenomena involved in the acquisition and interpretation of characterization results to convert them into a PSD. Also, there are new indirect approaches based on spectroscopy measurements. Even though the spectroscopy is a part of direct characterization techniques, the collected data from the spectroscopy images can be used in mathematical expressions for membrane characterizations. They can be called spectroscopy-based indirect approaches and they are preferable methods for the sub-nanometer scale membrane characterization while the well-established equation-based indirect approaches cannot characterize the sub-nanometer scale membranes. Accordingly, we can enumerate the following techniques:

a)	Porometries based on Young-Laplace Equation: among those here, we will consider the Liquid Displacement Porometries (LDP) and the Mercury Intrusion Porosimetry (HgP).

b)	Porometries based on Kelvin Equation: Evapoporometry (EP), Permporometry (PmP), and Gas Adsorption-Desorption (GAD) will be analysed in this group.

c)	Porometries based on Gibbs-Thomson Equation: here Thermoporometry (ThP) will be discussed.

d)	Porosimetries based on spectroscopy-related indirect approaches: Positron Annihilation Lifetime Spectroscopy (PALS) and Synchrotron Radiation (SR) will be investigated.

Indirect methods have a mathematical background, can observe a wide membrane surface area, and less expensive than the direct methods. Fig. 1 presents the classification of membrane pore size characterization techniques that are reviewed in this study.

Fig. 1: Membrane pore size characterization techniques that are reviewed in this study.

The projected techniques in Fig. 1 have their own procedures which can depend on the membrane type and application details. Hence, there are different review studies that have been presented to summarize and better understand the membrane characterization techniques in various applications. One of the oldest review studies were presented by Kim et al.[25]. Six different ultrafiltration membranes, which are either polysulfone or cellulose membranes, were characterized by using the thermoporometry, permporometry, and molecular weight cut-off experiments. In that concept, the review of mathematical backgrounds, operating conditions, and experimental setups were done. Another review study was performed by Nakao[5]. The characterization techniques were classified as the microscopic methods, bubble pressure and gas transport method, mercury porosimetry, permporometry, and thermoporometry. A detailed literature study was done, and the theoretical backgrounds were presented with the experimental procedure and evaluation of the results. At the beginning of 2000, Zhao et al. [26] updated the literature review study only for the hollow-fiber membrane characterization under the concepts of microscopic techniques, thermoporometry, water permeability, and molecular gas transport models. Recently, Rahman et al. [27] summarized the indirect membrane pore size characterization techniques only for ceramic membrane types. The experimental procedures and results of the gas adsorption-desorption, themroporometry, permporometry, bubble point, liquid displacement method, and mercury porosimetry techniques were explained and compared to one another. Several other reviews and books on membrane characterization have been published over the years, from more recent papers [28-30] to books [31], or book chapters [4]. All the foregoing review studies enlightened the researchers about the membrane pore size characterization techniques. They made a general revision of all techniques leading to membrane characterization, in both structural and performance-related aspects, or devoted to techniques leading to membrane surface knowledge [29-33].

Following the outcomes and discussions of the foregoing review studies, this study aims to present an updated, comprehensive, and more detailed review work for the indirect membrane pore size characterization methods. The current study brings some original perspectives that are explained as follows: it focuses on the indirect methods that are illustrated in Fig. 1 by collecting their initial studies from the past to the present. Moreover, along these years more research works have been presented related to the conventional characterization methods and even, new methods of promising features as Evapoporometry (EP) have been developed [34]. Therefore, an updated review of the techniques which can give information about the PSD of synthetic membranes seems necessary. Also, this study gives the theoretical backgrounds of the indirect methods with their derivations; hence, a better understanding of the relation between the experimental procedure and mathematical background is aimed. Besides, the spectroscopy-based indirect methods are reviewed for the membrane characterization of sub-nanometer scale membranes such as tight NF and RO membranes. Since the spectroscopy-based indirect approaches are relatively new compared to the well-established indirect approaches, they have not been reviewed in detail up to now, as far as we know. Along with this manuscript, we aim to review all these indirect properties, through the analysis of their assumptions and basis, the consideration of their different operational modes and the commercial setups available for each technique. Finally, selected results showing the features, potentialities, and drawbacks of each technique will be commented on. In the following pages, the abovementioned theories, techniques, and studies are presented in a sequence; then, an overall discussion is given.

# Indirect Method Theories

# Young-Laplace Equation

The Young–Laplace equation, which explains the pressure difference across the interface between two static fluids, is of fundamental importance in the study of capillary phenomena. In its more general form, the Young-Laplace equation expresses the pressure difference (ΔP) at the interface between two immiscible fluids for a general geometry as:

ΔP = 2γH        						(1)

being γ the interfacial surface tension of the immiscible fluids and H the interface curvature. Based on geometric considerations, this curvature can be obtained from two main curvature radii as:

H=121r1+1r2					(2)

where r1 and r2 are the principal radii of curvatures for the interface.  For the case of a cylindrical tube (i.e., the ideal membrane pore), the Young-Laplace equation must be modified to include the contact angle (θ) between the spherical meniscus formed in the inner pore and the pore wall, such that the pressure differential ΔP for a pore of diameter d is given by:

ΔP=4γcosθd			 	                  (3)

where γ is the interfacial surface tension of the immiscible liquids, and θ is the contact angle between the interface and the pore wall. For the case of a membrane, which is made up of a myriad of pores that can be simplified as straight-through cylinders with a range of diameters, Eq. (3) can be applied to determine the PSD by measuring the range of ΔP values [9].

# Kelvin Equation

The Kelvin equation is used in various indirect membrane characterization techniques such as evapoporometry (EP), permporometry, and gas adsorption-desorption. The derivation of the Kelvin equation accounts for the effect of surface curvature on the thermodynamic equilibrium of a pure liquid. As illustrated in Fig.2, the interface between the liquid and its vapor phase can either be planar (Fig. 2a) or a positive curvature (Fig. 2b).

Fig. 2: Thermodynamic equilibrium of a pure liquid with its vapor phase: (a) planar interface, and (b) positive curvature interface.

The chemical potential of a liquid and its vapor phase must be equal to each other at thermodynamic equilibrium, as defined in Eq. (4):

dμL=dμV 						 (4)

where μ denotes chemical potential, and the subscripts L and V represent the liquid and vapor phases. The liquid and vapor phases have the same total pressure at P0, but a positive curvature (Fig. 2b) means that the liquid phase has a higher pressure (PL) than the vapor phase (PV). For a planar interface (Fig. 2a), incorporating the ideal gas law gives:

VVdP=VLdP ⇒ VLdP=RTdPP				(5)

where molar volume is denoted by V. By assuming constant liquid density, integrating Eq. (5) allows us to relate the pressure for a positive curvature (Fig. 2b) to that for a plane (Fig. 2a):

VLPL-P0=RTlnPVP0 					(6)

Re-arrangement of Eq. (6) gives:

VLPL-PV-VLP0-PV=RTlnPVP0 				(7)

Invoking the Young-Laplace equation:

PL-PV = 1r1+1r2γcosθ				(8)

where r1 and r2 are the principal radii of curvature of the curved interface, θ is the liquid contact angle in a pore (at the gas-liquid interface), and γ is the surface tension. Assuming r1=r2=r for a hemispherical interface in a pore, Eq. (7) can be re-expressed as

2VLγcosθr-VLP0-PV=RTlnPVP0				(9)

Neglecting the second term on the left-hand side, because it is significantly smaller than the first term, gives the Kelvin equation:

lnPVP0=2VLγcosθRTr						(10)

If the liquid phase has negative curvature (as opposed to the positive one in Fig. 2b), the right-hand side of the Kelvin equation (Eq. (10)) takes on a negative sign:

lnPVP0=-2VLγcosθRTr					(11)

For the EP method for membrane PSD characterization, the Kelvin equation is adapted as follows:

lnPAPA0=-4VγcosθdRT					(12)

where PA is the vapor pressure of the volatile wetting liquid, PA0 is the normal vapor pressure of the volatile wetting liquid, V is the molar volume of liquid, and T is the absolute temperature during the EP experiments. According to Eq. (12), the vapour pressure of the liquid is related to the diameter of the pores, and therefore the pore diameters (d) of the membrane can be calculated if the range of PA values are known. The ratio of vapour pressures in Eq. (9) is equivalent to the ratio of the evaporation rates, since the driving force for evaporation from the pores is the differential vapour pressures. Hence, instead of PA, Eq. (12) can be defined by the evaporation rate above a flat liquid layer without a vapour-pressure depression, WA, as shown in the following:

d=-4VγcosθRTlnWAWA0						(13)

The evaporation rate, WA, can be determined via an instant mass change in the test cell. Placement of the test cell on a load cell or a high-resolution microbalance permits this determination, as per the evapoporometry (EP) method. The Kelvin equation underlies the evapoporometry, permporometry, and gas adsorption-desorption methods.

# Gibbs-Thomson Equation

The Gibbs-Thomson equation, which underlies the thermoporometry technique for characterizing membrane PSD, describes the thermodynamic equilibrium conditions of the solid, liquid, and gas phases of a pure substance inside the porous medium. The liquid amount in the porous material strictly depends on the pore size and inherently affects the interfaces during phase transformation. Brun et al. [35] presented a detailed derivation by using a combination of theoretical, numerical and experimental approaches. To develop the Gibbs-Thomson Equation-based means for pore size characterization, all the phases of a substance are denoted by the subscripts of i, j, and k. The Gibbs-Duhem equations are expressed as follows for the phases and interphases, respectively:

SidT-VidPi+midμi=0					(14)

SijdT-Aijdγij+mijdμij=0					(15)

where S is entropy, T is temperature, μ is chemical potential, m is mass, γ is interfacial tension, V is volume, and P is pressure. It should be noted that the second term of Eq. (15) is obtained via the Laplace equation that defines the pressure difference with respect to the interfacial tension, area, and volume:

Pj-Pi=γijdAijdVj						(16)

At the thermodynamic equilibrium of a pure substance at the triple point, all the chemical potentials (phases and interphases) become equal to one another, as expressed in the following differential equation:

sS-sVvS-vV-sL-sSvL-vSdT=VVvV-vSdγSVdASVdVV-VLvL-vSdγSLdASLdVL			(17)

where the solid, liquid, and gas phases are denoted by the subscripts S, L, and V, respectively. According to Eq. (16), the term dAijdVj, which can be defined as the curvature of two of the three interphases, must be known to find the temperature at the triple point. This is challenging for a porous medium[35], but it can be assessed in the region where the gas-solid interphase is planar, whereby the term dASVdVV approaches zero and vV≫vL. Also, the term sS-sL can be defined as the solidification entropy of the condensate, ∆Sf, and thereby Eq. (17) can be re-written as follows:

∆SfdT=VLdγSLdASLdVL=0				(18)

The term dASLdVL is defined as -2Rn, where Rn is the curvature radius of the solid-liquid interphase by assuming spherical interphase. Hereby, we obtain:

1Rn=12γSLT0T∆SfVLdT					(19)

The solidification entropy of the condensate, ∆Sf, is not a measurable term, but it can be expressed via thermodynamic state functions. Assuming incompressible flow, ∆Sf is defined as[35]:

∆Sf=∆So+T0TcS-cLTdT+∂vL∂TP-∂vS∂TPPS-P0+∂vL∂TPPL-PS	(20)

where ∆So is the normal solidification entropy, c the specific heat capacity at constant pressure, PS the vapor pressure of the undivided solid at the corresponding temperature, T. Following the Laplace equation (Eq. (16)), Eq. (19) and Eq. (20), the relationship between Rn and temperature can be defined as:

PS-PL=2γSLRn=T0T∆SfVLdT				(21)

Eq. (18) simply deduces that the relationship between the curvature radius of the solid-liquid interphase and the temperature can be explained if the solidification entropy of the condensate can be solved. According to the abovementioned theoretical development, the term ∆Sf can be calculated for liquids with respect to their thermophysical properties that define the state functions (see Eq. (17))[35]. As for the solid-liquid interphase, the curvature radius is also called the radius of the critical nucleus for the applied liquids. To determine  Rn, it is required to find the peripherical layer thickness of the condensate that does not undergo a change of state. For this purpose, a calorimeter is used and the required values of the terms γSL and Rn can be obtained for the corresponding liquid. In addition to γSL and Rn, the pore radii, Rp, can also be determined. By this way, the term Rp can be measured via the triple point temperature depressions at the solidification and fusion cases. It must be noted that these measurements are unique for the liquids. In the thermoporometry methods, which use the foregoing theoretical explanation, the temperature depressions are observed by using thermograms at both solidification and fusion cases. Following the thermogram observations, calorimetry measurement, and the explained theory, the relationship between Rp and temperature can be expressed in Eq. (22). Similarly, the equation of the distribution curve is given in Eq. (23),

Rp=-A∆T+B						(22)

∆V∆Rp=k∆T2Way						(23)

where A and B are the constants obtained from the numerical solutions (e.g., A=64.67 and B=0.57 for water [35]), k the sensitivity of the calorimeter setup, and Wa is the apparent energy (e.g., solidification energy).

# Spectroscopy-Based Theories

# Positron Annihilation Lifetime Spectroscopy (PALS)

Positron Annihilation Spectroscopy (PAS) is a direct observation method that can be used for membrane pore size characterization even though it was firstly developed for the electronic structures [36, 37]. However, the collected images via the PAS can also be used for the analytical modelling studies of the membrane pore size characterization. Compared to the well-known direct observation methods like TEM and SEM, the PAS can better determine the subnanometer-scale membrane holes; thus, it is preferred for the analytical modelling of the reverse osmosis (RO) and tight NF membranes. When PAS is used to analyse pore or holes inside a membrane or other porous materials, the resulting technique  is usually named as Positron Annihilation Lifetime Spectroscopy (PALS) technique [36]. By means of the advantage of the PAS in small-scale membrane holes, the PALS achieves more reliable pore size characterization than the other established indirect methods such as GAD and similar differential scanning calorimeter approaches [30, 38]. In simple words, the PAS method observes the active skin layer of the RO/tight NF membranes, and the obtained data are evaluated by the analytical model of the PALS. It measures the lifetime spectrum of positrons; then, the pore radius is determined on the active skin layer of the membrane [39, 40]. Below, the theoretical backgrounds of these two methods are presented.

In the PALS analysis, the positrons are injected to the RO membrane sample and they are annihilated with the electrons to two γ rays, or combined with the electrons that form a positronium (Ps) atom with a vacuum binding energy of ~ 6.8 eV [41]. Two different Ps states exist as the antiparallel-spun (para-Ps or p-Ps) and the parallel-spun (ortho-Ps or o-Ps) with respect to the spinning positioning of the positrons and electrons. In the p-Ps state, the total spin is 0 (zero) and the intrinsic lifetime (τp-Ps) is roundly 0.125 ns with two- γ decays while the total spin is equal to 1 (one) and the intrinsic lifetime (τo-Ps) is up to 142 ns with three or more γ decays. In polymeric membranes, the τo-Ps is 1 to 5 ns since the positron is annihilated with one of the surrounding electrons in an opposite spin [36, 41]. The free-volume holes in the membrane can be detected by the PAS data because the annihilation photons move in these open spaces. Refs. [42, 43] correlates the measured τo-Ps with the free-volume hole radius, rfv, as shown in Eq. (24),

τo-Ps=λ1-rfvrfv+Δrfv+12πsin2πrfvrfv+Δrfv-1				(24)

where λ=2 ns-1 is the annihilation rate of spin-averaged Ps and Δrfv=1.656 Å3 is the valid homogeneous layer for polymers [41]. As within Like other indirect methods, the holes are assumed infinite spherical structures so that the radius can be defined as rfv-0=rfv+Δrfv. For this assumption, the volume of the spherical structure can be written as in Eq. (25),

VF=4πrfv33							(25)

The PALS measurement achieves very sensitive results, but it must be noted that the operating conditions should be maintained well for accurate results because the fluctuations of the operating parameters can negatively affect the measurements and may result in uncertainties [44]. The simplified experimental setup schematic of the PALS method is shown in Fig. 3.

Fig. 3: Simplified experimental schematic of the PALS technique (redrawn and modified from Ref. [36]).

# Synchrotron Radiation (SR)

Synchrotron radiation (SR) is the electromagnetic radiation produced by a synchrotron storage ring that accelerates charged particles nearly to the speed of light when those charged particles, relativistic electrons, are expedited in a circular path by strong magnetic fields [45, 46]. SR sources provide opportunities for real-time studies thanks to their spectral properties such as continuous wavelength spectrum generation and luminosities production with a large range of spectrum [47]. Therefore, it is a preferable technique that provides chemical imaging analysis with high-resolution [48]. The images (radiographs) collected via the SR method can be used for further membrane pore size characterization, in the same indirect manner as other membrane microscopic images.

# Young-Laplace Equation-Based Techniques

# Liquid Displacement Porosimetry (LDP)

Liquid displacement porosimetry (LDP) consists of gas-liquid displacement porosimetry (GLDP) and liquid-liquid displacement porosimetry (LLDP) techniques. Therefore both gas-liquid displacement porosimetry (GLDP) and liquid-liquid displacement porosimetry (LLDP) will be explained in the next sections; but their pioneering approach, the bubble point test will be presented first.

# Bubble-Point Test

As the first technique based on the Young-Laplace equation, the Bubble Point method was used to analyse the size of the pores present in a filter [49]. The method is now being used especially by membrane manufacturers to assure the integrity of their products. Several devices have been developed by the main membrane companies which commercialized them as Integrity Tests, for example, Integritest 5 (Millipore), BP Tester (PMI), Flowstar IV (Palltronic) or Sartocheck 5 (Sartorius) [50]. They are based on the primal idea of Bechhold [49].

The chief problem with using this technique as the integrity test is the relatively high surface tension of the air-liquid interface. In the original work by Bechhold [49], water was used as an impregnating fluid. Since the water-air interface has a surface tension of 72.8 mN/m, a pressure of around 15 bar is needed to analyse a membrane with a maximum pore size of 0.2 mm. Later on, researchers found better wetting fluids, but even using typical halogenated fluids to wet the membrane (which assure almost perfect membrane wetting with low contact angles that, for practical purposes, can be taken as zero), the surface tension with air is around 16 mN/m, which means that analysing a membrane having a maximum pore size of 5 nm requires a pressure of 64 bar. Such high pressures could lead to membrane structure distortion, compromising the feature that the technique is supposed to characterize. Because of the high pressures associated with small pore sizes, the Bubble Point Test is commonly used for MF and UF membranes.

According to the Young-Laplace equation, the corresponding pressure for first air bubbles appearing gives the size of the biggest pore present in the membrane if the pressure applied to the air is increased steeply. Obviously, further increases in air pressure do not serve to give information about smaller pores opened, since new air bubbles are not distinguished from initial ones. Therefore, the method cannot be considered suitable to obtain the pore size distribution of a membrane, except in the very improbable case of a completely homoporous membrane (i.e., all the pores are of exactly equal sizes). Even though the method has not been recently preferred for the membrane pore size characterization since the LLDP and GLDP techniques seem more efficient and attractive, some of the bubble-point methods are worth to mention. Reichelt [51, 52] applied the bubble point technique to characterize larger membrane sample areas. The method successfully tested the cartridges with polypropylene capillaries and flat polyamide membranes which have filtration areas from 0.5 m2 to 6 m2 [53]. The method is also used to characterize the pore size distribution of various geotextile samples [54, 55]. Hernandez et al. [56] used the bubble point test to characterize several track-etched microporous polycarbonate membranes with nominal pore sizes ranging between 0.1-5.0 µm. Jakobs and Koros [57] characterized two different types of ceramic membranes which were flat sheet and tubular alumina membranes via bubble test method. Yu et al. [58] modified the conventional bubble point method to measure the narrowest area at a pore tunnel (pore-throat) of membranes. Nassehi et al. [59] developed a numerical method to analyze the bubble point test by considering the surface tension effect at the gas-liquid interface and validated their method with experimental results of track-etched polycarbonate membranes. The bubble point method is also used for the characterization of supported ionic liquid membranes by using different porous ﬂat sheet supports [60].

# Gas-Liquid Displacement Porosimetry (GLDP)

The original idea on bubble-point [49] only served for the evaluation of the maximum pore size present in a filter, but the technique can be easily extended to get information about the whole pore size distribution (PSD) through increasing further the pressure after the first appearing bubbles. As the pressure progressively increased beyond the bubble point, successive pores of decreasing sizes gradually empty, giving rise to the GLDP technique, which is a combination of the bubble point test and solvent permeability method to obtain the PSD of the membrane [61]. The method was later improved to take into account the surface energy at the liquid-gas interface [62, 63]. In effect, GLDP includes the measurement of the gas flow through the membrane pores to get information on the PSD. To convert the gas flow increments into the contribution of newly opened pores, a transport model is needed. A simple choice could be to use the Hagen-Poiseuille law for convective flow inside capillaries, mostly valid at the typical applied pressures and pore sizes involved. However, this model predicts a linear relationship between flow and pressure in all the ranges, while experimentally some non-linearity of the flow curve as pressure increases can be observed. This is because the Knudsen diffusion model should be contemplated for pores where the mean free path of the gas molecules, , is much lower than the capillary diameter. In fact, calculations usually consider some sort of combination between both models of flow. Schneider and Uchytil [64] showed the discrepancy between GLDP results and real pores due to the improper consideration of both gas transport models involved. Hernández et al. [65] proposed how to solve such discrepancies through a factor leading to a smooth transition from Knudsen diffusive flow to Poiseuille convective one. With such an approach, GLDP calculation resulted in a much more accurate agreement with actual PSD and porosities of test membranes. Another approach by Shao et al. [66], who proposed that the contribution of Knudsen diffusion in GLDP can be effectively suppressed by an appropriate selection of the pore-filling liquid. Departing from the simple idea of Erbe [61] and using different modeling for PSD calculation, several commercial types of equipment were developed, including automatically controlled data acquisition and all subsequent calculations of PSD.

The first commercial equipment appearing in the market was the Coulter Porometer, but this successful equipment (especially the improved version, Coulter Porometer II) shortly ended fabrication and Beckman, which acquired the porometry division of Coulter, discontinued its production. Among other GLDP devices, that of PMI (Advanced Capillary Flow Porometer its latest version) deserves a commentary, since it is one of the first in the market that claims to characterize pores down to almost the nanometer range. Other companies manufacturing air-liquid porometers are Quantachrome (featuring Porometer 3G), Porometer (a company originated at Benelux Scientific which features the Porolux 1000 as the latest version), PSM165 (TOPAS) or PRM-8410 GLDP (IFTS). From an experimental point of view, all types of GLDP equipment in the market use a quite similar approach. The first step involves filling the membrane to be analyzed with a liquid that enters all the pores. Then, the membrane is placed in a dead-end cell and subjected to increasing pressures by controlled air supply. As pressure increases, it reaches a value corresponding to the bubble point of the Bechhold method, when the biggest pores in the membrane are opened and the air starts to flow through. As pressure continues to build up steeply, pores of decreasing sizes are progressively emptied and the flow of air through the membrane increases. A flow meter placed downstream allows determining the successive gas flow increments. The pressure applied and the size of the pores opened at each pressure are related through the Young-Laplace equation, Eq. (3), whereby the contact angle at the air-liquid-membrane interface plays a key role. The contact angle depends on membrane-liquid interaction, and therefore different wetting liquids with appropriate hydrophilicity/hydrophobicity characteristics should be used for analyzing different membranes. However, in practice, the choice of liquid is very complicated and risky, as very often the membrane characteristics are not known. Therefore, generally, a more or less standardized liquid, which exhibits both hydrophobic and hydrophilic radicals and thus able to wet reasonably well all kinds of membrane materials, is used. Other desirable features of wetting liquids include low surface tension, low vapor pressure, and low reactivity. Perfluoro halogenated compounds are the most successful liquids on meeting such requirements and accordingly most GLDP manufacturers usually supply these liquids (Porofil®, Silwick®, Porewick®, Galwick®, those latest based on different configurations of Fluorinert®) to their customers [65]. A plot of airflow downstream of the membrane versus applied pressure is usually called a wet flow curve. Most GLDP equipments perform a second run starting from the initial pressure and repeating again all pressure steps. Once more airflow values are recorded, the plot of these new values results in the dry flow curve. From the comparison of the PSD’s of both curves, other important related parameters such as porosity, mean pore size, mean flow pore, or estimations of membrane permeability are obtained. A combination of convective flow (Hagen-Poiseuille) and diffusive flow (Knudsen) models are used for converting flow-pressure information into the contribution of each pore size to the total permeability [65].

After the appearance in the market of the automated setups, the GLDP technique has been extensively applied for membrane characterization. Thanks to its capabilities, not only can synthetic membranes be studied through GLDP, but it can also be used to obtain porosity information for many other porous materials, although only information about pores traversing both sample sides (i.e., only pores that contribute to flow) can be obtained. The upper pore size limit can be extended to 200 microns, which makes this method possible to use to characterize all kinds of sieving filters along with woven and non-woven textiles [4]. Regarding the lower limit of applicability, it depends strongly on the properties of the wetting liquid (namely, both contact angle and surface tension), but for most liquids and commercial porosimeters, the technique can be used down to 50 nm by using applied pressures of 14-15 bar. Some companies offer GLDP porosimeters that are able to measure pore sizes as small as 13 nm, but it is worth noting that pressures as high as 34 bar would be needed, which increases the risk of membrane failure and structure distortion. For such smaller pores (roughly under 50 nm), LLDP (which will be detailed in the next section) offers advantages in terms of lower applied pressure and accuracy that should make it more practical compared to GLDP. As a reliable and robust method for the characterization of MF membranes, fibers, and non-woven textiles, GLDP has become a recommended standard [67, 68]. It has been used to obtain PSDs for many research or commercial membranes [69], including flat-sheet [70], nanofibrous [71, 72], and hollow fiber [73] membranes. A number of researchers used GLDP to characterize ceramic membranes, including the maximum pore size of porous Ti3SiC2 [74], the mean pore size of the Al2O3 tube substrate coated with pencil lead and carbon membrane [75], and pore-throat size [58].  Also, in the literature, some patents related to different applications of GLDP can be found [53, 76]. A recent work by Mourhatch et al. used porometry to determine the pore size distribution of silicon carbides, approximately modelled as a bundle of capillaries [77]. The technique has been also used to study membrane fouling [78-80], accounting for the decrease of mean pore size and shifting of PSD due to foulant deposition inside the pores. As a reference technique for porous membrane characterization, it has been quite often used to evaluate pore size characteristics of membranes compared to different techniques (e.g. HgP, SEM, porometry, image analysis), and the results seem in good agreement [81]. Also, recent papers continue to model the gas flow to increase the accuracy of pore number calculations [82, 83]. The GLDP has been also used to test the accuracy of newly proposed characterization methods due to its recognized reliability [84]. In the literature, the technique is sometimes named Capillary Flow Porometry [72], Liquid Extrusion Porosimetry [85], or Flow Permporometry [81], but all of them refer to the same procedure, or slight variations of it. Most devices use some form of a data-smoothing algorithm for the typically more than 200 pressure steps in a complete analysis to result in approximately Gaussian PSDs.

The Liquid Extrusion Porometry is a variation of the GLDP technique based on the contact between the sample to be analysed and a membrane having pores much smaller than the sample, known as capillary barrier membrane, and the extrusion of wetting fluid from the sample [86]. Sample and capillary barrier membrane are completely immersed into the wetting fluid.

Jen and Gupta, developed for PMI, the method to be applied for different porous samples [87]. This setup from PMI has been used by Manickam and McCutcheon [85], to study several types of non-woven MF membranes. The method can be simplified also to determine the Liquid Entry Pressure (LEP) [88], and the PSD of hydrophobic samples [89]. The method has been used recently to determine the capillary pressure, i.e.  the difference between the water and gas pressures at equilibrium. The measurement is based on determining the difference in the pressures to be applied during water injection and withdrawal for polymer electrolyte membrane fuel cells (PEMFC), being this capillary pressure an important parameter to assess the performance of a fuel cell [90]. The microstructure of a porous medium is generally quite different from a simple bundle of parallel cylindrical tubes. Most porous systems are rather characterized by a system of highly interconnected pores. In fact that the use of the normalized flow-rate curve  in conjunction with the capillary pressure curve always lead to a good estimate of the PSD for mostly heteroporous membranes.

The experimental procedure is very similar to GLDP, the sample is first saturated in a highly wetting fluid, which is a fluid with low surface tension, low viscosity and vapour pressure. Then the capillary pressure, for a non-wetting fluid (normally air), is slowly increased and the wetting liquid flows out from the sample onto an analytical balance. As the system reaches equilibrium at each capillary pressure step, the mass of liquid drained is recorded and the pressure is again increase. Once the maximum pressure has expelled the wetting liquid from all the pores of the sample, i.e. there is no contribution of the liquid drained, the capillary pressure is slowly relieved and the wetting liquid flows back into the sample. Using these draining and flowing back data, the capillary pressure curves (showing frequently some hysteresis) can be generated [125]. IFTS has designed a commercial setup to accomplish such LEP adapted to fuel cell characterization.

Summarizing all comments about GLDP (and variations of that) it can be stated that the technique offers some advantages over other PSD characterization methods, including (i) it is non-destructive [81]; (ii) as a flow-based PSD technique, it determines the smallest diameter in continuous pores that are open to flow; and (iii) it requires minimal sample preparation, small footprint and simple procedure [9]. On the other hand, drawbacks include: (i) indirect method and thereby it requires justified assumptions on pore geometry; (ii) limited accuracy in measuring pressure and flow rates; (iii) problematic for low-porosity samples; (iv) PSD compromised by interconnecting pores and affected by choice of wetting liquid [91]; and (v) inability to characterize biofouling layer and difficulty to correct for adsorbed t-layer.

A quite common criticism against GLDP and related techniques is that the method cannot distinguish between different inner pore geometries [81], but only determines the narrowest section across the pore (i.e., throat diameter). Certainly, this criticism is true and therefore the technique is limited in properly analyzing inter-connected pores. But this drawback is not as critical for membranes, whereby pores act as sieves (which is very often the case), this smallest pore cross-section dictates the feed constituents that pass through or get retained by the membrane and consequently determines the performance in terms of separation extent and selectivity. Another important criticism against liquid displacement porometries, raised against both GLDP and LLDP, concerns the assumption of certain geometry to interpret and convert measured data into PSD. Effectively, when using any transport model (Knudsen, Poiseuille, or a combination of both), generally it is assumed that pores are perfect cylindrical capillaries crossing normally the membrane from one side to another. However, the cylindrical geometry is only true for a few test membranes, while more complicated geometries exist for most membranes. This problem has been avoided sometimes by using different kinds of structural parameters, with a more or less empirical basis, to account for the differences between the actual geometry and the cylindrical one. These issues rely on the usual criticism attributed to all the Washburn equation-based techniques, which define the differential capillary fluid flow rate in a cylindrical tube. To have a clearer idea about the pertinence of such an assumption, some comments are warranted. Notably, pores are often so irregularly shaped that they cannot be considered any sort of tube but are rather tortuous pathways that remain open between the complex structures formed during the membrane manufacturing process. These pathways, along with cross-sections clearly differing from circular ones, exhibit significant variability of the diameter so it is almost impossible to mimic any portion of any ideal tube. Since the Laplace equation gives the highest pressure that is needed to intrude the fluid inside the porous structure, the pressure corresponds to the pressure required to pass the narrowest section of the pore. Even for non-tube-shaped pores, the cross-sections of some portion of the pore, particularly around the neck, can be considered more or less constant (see Fig. 4). The Washburn equation refers to such cross-sections, and so gives information about the minimum section found inside the pore, which governs the flux and possible separation properties of that pore.

Fig. 4: A simplified cross-sectional view of a non-cylindrical-shaped pore structure.

Regarding the shape of this cross-section, it is expected to be very different from a circular one, so that the Washburn equation again loses validity. Nevertheless, if we consider the original Young-Laplace equation, even an irregularly shaped meniscus should have a curvature that could be defined from the principal curvature radii (R1 and R2). Those radii are related to the dimensions of the cross-section of the pore, with one of them corresponding to the maximum dimension and the other to the minimum dimension found in that section. These could be length and width for a rectangular section or major and minor axes for an ellipsoidal one. When we substitute the complete Young-Laplace equation for a simplified one, we are considering that we move from an irregular shape to a circular one, in such a way that both have the same area. The error made in such approximation can be more or less evaluated. In effect, if Rd is between both extreme curvature radii, then we have:

R1≈Rd-ΔR

R2≈Rd+ΔR 	                                             (26)

So, the corresponding term in the precise Young-Laplace equation is:

1R1+1R2≈1Rd-ΔR+1Rd+ΔR=1Rd11-ΔRRd+11+ΔRRd		(27)

where the incremental term, ∆R/Rd, could be very small. This expression can be extended as a sum of power factors as follows:

11-ΔRRd=1+ΔRRd+12ΔRRd2+...

11+ΔRRd=1-ΔRRd+12ΔRRd2-…

(28)

Which leads to:

1R1+1R2≈2Rd+2RdΔRRd2 					(29)

The second term on the right-hand side can be neglected, supposing that (∆R/Rd) << 1. This assumption can easily hold true for most of the typical pore geometries, except for slit-shaped pores in which R1 = Rd, and R2 → ∞. So, the Young-Laplace equation becomes:

Δp=γRd						(30)

Which is the circular geometry that usually assumed. But this kind of pores are not so frequently found, and anyway, they can be analysed in a different manner, if their shape fits that one. In that way, what we are obtaining when using the simplified Laplace equation is a pore size that somehow averages the extreme real dimensions of the actual pores. This can be easily assumed as a convenient standard while leaving to researchers the effort to make more accurate approximations when needed.

As mentioned, there is a source of error in the determination of pore size distributions from the above-considered techniques. This appears when converting directly measured data for each technique (differential flows for both gas-liquid and liquid-liquid porometries and differential intruded volumes for mercury porosimetry). Usually, the differential distributions experimentally obtained in the commercial devices using each technique are subsequently converted into distributions (relative or absolute) of pore numbers, which involves a suitable model for the pore geometry in such a way that each flow or volume is converted into the corresponding number of pores contributing to that differential value. As already mentioned, this step is commonly accomplished by using a capillary pore model for the pores or even considering the Carman-Kozeny equation for granular structures or other convenient transport models. There is no need to explain again the errors that this procedure can give, as far as it would convert the technique in question into a model-dependent source of information. In that sense, a reasonable idea, at least for standardization procedures, could be to give these distributions as directly measured values, leaving to the user the interpretation of the actual shape of the pores which can be then easily converted in absolute pore numbers. Correspondingly, most pore size distributions could be compared directly using both flow and volume differential distributions, giving the range (maximum, minimum, and mean pore size) along with the shape of the distribution (skewed, log-normal or Gaussian distributions) can be for ready comparison among membranes.

Finally, another source of error found in the interpretation of GLDP results relies on the need to know the actual pore length to calculate the absolute number of pores of each size. This length is usually identified with the membrane thickness for symmetric membranes and active layer thickness for asymmetric ones. In either case, an assumption is made that the pores are normal to the membrane surface with no tortuosity. But the actual thickness is challenging to evaluate (especially for asymmetric membranes), and furthermore the pore length cannot be accurately quantified by filter thickness in view of tortuosity. Recent work has improved data interpretation by determining the tortuosity factor for each pore class in the distribution [92]. Another important point of controversy refers to the contact angle to be used, which is a difficult-to-evaluate parameter that is not always known for each membrane material and structural configuration.

# Liquid-Liquid Displacement Porosimetry (LLDP)

After the reasonable success of the GLDP technique, particularly with commercial devices appearing in the market, the next natural step should have been to extend the technique to lower pore sizes. As commented in the previous section, the typical lower size range of GLDP-based porometers is approximately 50 nm, which restricts the application to MF membranes or very open UF ones. For analysing membranes with pores smaller than 50 nm (as usually found in typical UF or NF membranes), the air-liquid interface requires high pressure. To this end, Liquid–Liquid Displacement Porosimetry (LLDP) circumvents the drawback by making use of a liquid-liquid interface inside the pores. Since the transport mechanisms of liquids and gases are not the same, the approaches used to convert GLDP data into PSD are not valid for LLDP and must be adapted to account for the liquid transport features. In fact, the liquid transport, being convective at the pressure conditions used, needs only a simple Hagen-Poiseuille model for convective transport inside capillary tubes in most cases. Two different approaches to making such conversion can be used, both leading to similar results. The first one is based on a graphical (or numerical) calculation of the successive permeability increments, after which the Hagen-Poiseuille model is used to determine the number of pores needed to get such permeability increment as bigger pores are opened [93-95]. The second involves a differential method for continuous calculation that can be also adapted for a group of discrete data points [96]. Both methods lead to similar results as both are finally based on applying the convective transport modelling inside capillary pores. It is clear that LLDP shows many similarities with GLDP, and their advantages and drawbacks are similar. Recall that the only difference between the two techniques is the use of a gas-liquid-solid versus a liquid-liquid-solid interface, the latter of which results in a reduction of the surface tension of the corresponding interface. As a result, the size of the pores that can be detected at reasonable pressures is very different between the GLDP and LLDP.

The LLDP uses two immiscible liquids to perform the porosimetric analysis. From these two liquids, the one with a higher affinity with the membrane material is normally employed to wet the membrane (named as the wetting liquid, whose key role is to get the most complete possible wetting of the membrane by filling all the pores), while the other liquid (namely, the displacing one) with lower affinity for the membrane, is used to displace the wetting one. It is easily understood that the roles of these liquids can be switched depending on membrane characteristics [97]. The pore size of the membrane is calculated based on the applied pressure, which is a function of the interfacial tension between the two liquids. In this method, it is assumed that the pores are cylindrical and distributed non-uniformly. The pressure, ΔP, to be applied to push the displacing liquid inside a pore of a given size, which is filled by the wetting liquid, can be described by the Young-Laplace equation [9] (see Eqs. (1) to (3)). The mechanism of displacement is determined by the actual pressure, the liquid surface tension, and the pore size, while the value of contact angle θ is considered null given that the membranes are assumed to be completely wetted by the wetting liquid [97]. Similar to what happens at the gas-liquid interface, the displacement phase will not be expelled from any pore if the applied pressure is insufficient to override the interface surface tension for the biggest pore (more specifically, the narrowest section of the widest pore) present in the membrane. As shown in the case of ΔP<ΔPth in Fig. 5a, where ΔPth is the threshold pressure (or the bubble point per GLDP nomenclature, even though in this case there is no bubble appearing across the membrane, but instead drops of the displacing liquid surrounded by the wetting one and normally indistinguishable from each other), the displacing liquid is unable to displace the wetting liquid from the pores. When the pressure is increased to ΔPth (Fig. 5b), the displacing phase can overcome the minimum surface tension and start to penetrate through the membrane (i.e., expel the wetting phase) via the largest pores (dmax). At this pressure, the flux of the permeating liquid permeating through the pore with radius rth will be Jth. The pore with radius r<rth will be permeated by the displacing liquid only if ΔP>ΔPth (Fig. 5c) and any further pressure increase will give a proportional increase of the flux across the membrane.

Fig. 5: The gas-liquid interface for different membrane pore sizes at different pressures (ΔP): (a) ΔP<ΔPth (b) ΔP=ΔPth, and (c) ΔP>ΔPth [97]. ΔPth is the threshold pressure, or the bubble point per GLDP nomenclature, even though in this case there is no bubble appearing across the membrane, but instead drops of the displacing liquid surrounded by the wetting one and normally indistinguishable from each other.

For each increasing pressure step, the corresponding flux (Jth) is measured to obtain the curve of flux versus pressure (Fig. 6), and thereby pore radius rth values represented by the permeability of the membrane (i.e., flow: pressure ratio) are obtained [9]. Therefore, by measuring the equilibrium pressure drop corresponding to each increment of flux, a pore size distribution of the membrane can be evaluated in terms of the contribution of each pore size opened to the whole membrane permeability (often termed as asymptotic permeability). This permeability corresponds to that of the whole membrane for the displacing liquid, at which all the pores are emptied of the wetting liquid. We can calculate the contribution of each pressure-flow step in the permeability distribution, and eventually the pore number distribution from the threshold step until asymptotic permeability (th = 1, ..., k) (Fig. 6). It is common to assume that the geometry assumption is not so restrictive since many pore geometries can be simplified to a bunch of more or less cylindrical pores [98] having a radius equal to the narrowest section of the actual pores found in the membrane structure.

Fig. 6: A representative pressure vs flux curve[9].

With micro-sized pores, the laminar flow should prevail and hence the water flow rate Qth drained from dmax can be described by Darcy’s law:

Jth=kthAth∆Pthɳl					(31)

where kth is the intrinsic permeability coefficient of the displacing liquid at dmax, Ath is the total area of the largest pores, η is the dynamic viscosity of the displacing fluid and l the pore length (which corresponds to the active layer thickness) [99].  For convective transport of liquids (see Fig. 5), the Hagen–Poiseuille model for the flow-through capillary pores can be used to derive the PSD:

Jk=Nk π  rk4Pk8ɳl					(32)

where rk is the pore diameter of the pores opened up in the kth step. For each increasing pressure step, ∆Pk, a corresponding volumetric flow, Jk, was measured.

The algorithm developed by Grabar and Nikitine [96] is essentially a differential algorithm, which requires the continuous curve of permeability variation for derivation. In this case, to directly use the obtained experimental data, the following equation is employed:

Nkd=8ηlπrk5∆Lk                                                     (33)

where Nk represents the number of pores of radius rk, i.e., in the k-th class per unit membrane surface; ΔLk is the ratio of successive flow and pressure increments between two experimental consecutive points (i.e., ΔLk= (Jk, − Jk-1)/(Δpi −Δpi −1)); and Lk corresponds to the k-th experimental permeability (i.e., Lk=Jk/Δpk ΔLk and Lk represent the pendant of the curve tangent to the porosimetric curve in the considered point, and the slope of the straight line passing from this point to the origin, respectively. Eqs. (3) and (30) are the fundamental basis of LLDP, from which further analysis can be performed to better understand the internal structure of the membrane[100].

The fluid displacement techniques have been considered, for years, as very promising solutions, because of advantages including non-destructive, very fast, and reveal only pores opened to flux (which are the only ones contributing to membrane permeation). It represented for many years the most reliable method to characterize PSD in the UF range. Nevertheless, LLDP has not found a wider utility due to the lack of effective commercial devices available in the market. The available ones include the LLP 1100A (PMI), Poroliq 1000 series (Porometer) and FFP PRM-8410 (IFTS). Due to limited availability, most researchers interested in using LLDP have to develop their own laboratory-based equipment designed and fabricated in-house. Among the dominant researchers focussed on LLDP, some are co-authors on the present review [6, 101-107]. Also, the group in Genoa, with which some of the authors developed a long-term collaboration, were very active on LLDP for many years [108-112]. These porosimeters used in LLDP analysis consist of an automated device developed in the laboratories of the Group of Surfaces and Porous Materials, SMAP, of the University of Valladolid and the group of Membrane Processes and Membrane Technology, MP&TM, of the department Chemistry and Industrial Chemistry of the University of Genoa. A detailed description of the equipment and experimental procedure can be found elsewhere (Fig. 7) [97]. The main feature of the equipment is the use of a precise syringe pump (ISCO-500D), allowing accurate and very stable fluxes through a dead-end membrane cell, without fluctuations and thus negating the need for any sort of dampening.

Fig. 7: Scheme of the LLDP dispositive from SMAP and MP&TM groups [97].

Both porometers are completely automatized based on laboratory developed software [107], which includes software-controlled operation, data acquisition, and data treatment running on the LabVIEW® platform.

Fig. 8: Schematic diagram of the IFTS porometer : (1) Air supply; (2) Air pressure tank;  (3) Feed line; (4) Membrane housing cell; (5) Disposal; (6) Analytical balance; (7) Retention line (pressure sensor); (8) IFTS LLDP software [9].

Another setup, originally intended for research but nowadays being marketed, is that of IFTS which is based on an early patent by Thierry Courtois [111]. This porometer consists of an automated and commercialized pressure-constant device developed by IFTS (Fig. 8) for the testing of pore size as small as 2 nm and uses lower pressures for characterization of porous membranes, allowing liquid permeation through polymeric membranes across very accurate and stable pressure. The fluxes are measured using an analytical balance (Sartorius, Practum, the accuracy  1 mg). This LLDP-based porometer is able to measure several structural and functional parameters of a porous material (including hollow fibers, tubular and flat sheet membranes), namely, mean pore diameter, pore size distribution, solvent permeability and droplet point. Liquid pairs with very low mutual interfacial tension and fully (or at least almost fully) immiscible are very suitable for testing membranes having pore sizes in the nanometer range since these pores can be invaded by the displacing liquid at very low pressures that do not alter the membrane structure. For example, according to Cantor equation (Eq. (3) with ; i.e. for perfect wetting (θ=0°)), for Water/Isobutanol system (γ=1.7 mN/m), the wetting liquid which fills a pore with a radius of 0.034 µm will be expelled out by applying to the displacing liquid pressure of 1 bar, i.e. a value very reasonable and for which membrane compaction (a factor representing an important source of error in LLDP measurements) [73] is negligible. A great advantage of this method (common to GLDP as stated before) lies in the fact that the membrane is tested under wet state, i.e. in a condition very close to that occurring in largely important membrane separation processes, such as dialysis, MF, UF, or NF. The method is clearly sensitive to interactions between the membrane and the liquid pairs. Swelling phenomena and membrane compaction (affecting especially skin layer) during testing may deeply alter the permeation through the pore thus leading to unreliable results. Membranes are prone to swell when the porous material (very often polymeric) is in contact with organic solvents due to the interaction between materials, comprising the membrane and molecules of the organic solvents, and then they will result compacted after finishing the analysis of the whole PSD [113]. This effect could be important on a proper determination of the actual pore size distribution of the membrane, which later should be useful in the selection of appropriate membranes for the effective separation process. Several liquid pairs (wetting-displacing fluids) can be selected to perform LLDP analysis. Those liquid pairs should be selected according to expected values of membrane porosity and pore size range [9]. For example, when analyzing a UF membrane, having most of the pores in the range from some to 50 nanometers, the isobutanol/water mixture (presenting a surface tension value of around 1.7 mN/m) is a reasonable and quite common selection. In any case, most of the main liquid pairs available are shown in Table 1, along with their main characteristics.

Surely the earliest LLDP study was performed by Erbe [61] which explored two possibilities: first using the initial air-water interface proposed by Bechhold, or also a mixture of two immiscible liquids (water and isobutyl alcohol, also proposed by Bechhold [49], for which he quoted value of 1,73 dyne/cm for its surface tension), then reducing the pressure needed to analyze ultrafilters and allowing to extend the original Bechhold bubble point method to the whole membrane PSD. Concerning the use of alcoholic mixtures in water (as most possibilities are shown in Table 1), Germic et al. [114] pointed out the possible problems of swelling effect due to the adsorption of alcohol inside the membrane material matrix. To minimize such problem Phillips and DiLeo [115] have proposed a liquid porosimetric CorrTest™ based in the use of a non-alcoholic containing mixture to perform LLDP analysis, method being developed by Millipore. They used two fluids formed upon mixing polyethylene glycol 8000 (PEG-8000), ammonium sulfate, and water in specific mass proportions. The resulting PEG-rich phase and ammonium sulfate-rich phase correspond to the wetting and intrusion fluids respectively. This biphasic system formed by mixing two polymers or a polymer and a salt in water is known as aqueous two-phase systems (ATPS) [116] and can be used for separation of cells, membranes, viruses, proteins, nucleic acids, and other biomolecules. The interfacial tension between the two aqueous phases plays a major role in the interfacial tension which between the phases can be 3−4 orders of magnitude smaller than conventional fluid-liquid systems :< 0.5−mN/m for ATPS compared to ∼72 mN/m for the water−vapour interface. Ultra-low surface tension through ATPS mixture strongly depends on the phase diagram of ATPS [117-121] and strongly depends on the molecular weight of the polymer so the interest to reduce droplet point pressure and no swelling have a great interest for LLDP. Normally, it involves the preparation of the stock solutions of all the phase components, which are then mixed in appropriate amounts (by weight), and then the emulsion is allowed to separate into two phases under gravity or in a centrifuge after complete phase separation is achieved [122]. We have proposed a two immiscible fluids answering the next targets: the first one is to reduce polymer swelling with very low interfacial tension; the second one is to get, according to with Physical-Chemical criteria, good chemical compatibility with Liquid Liquid Porometer especially internal parts to avoid corrosion or degradation and chemical membrane inertia; finally to have the properties of a Newtonian liquid with low displacement viscosity and low vapour pressure. Mixtures of polyethylene glycol 1000 (PEG-1000), Magnesium sulphate anhydrous and water in specific mass proportions have good perspectives in filling all these requisites.

In order to determine the interfacial tension, it was initially used the spinning drop technique [118, 122] of different samples based on mixtures of polyethylene glycol (PEG) with magnesium sulphate solutions. The spinning drop tensiometer technique allows measuring the interfacial tension between two liquids phases by analysing the dimensions of a drop in a rotating capillary by the Vonnegut method[123, 124]. It is a precise and simple measurement to determine extremely low interfacial tensions. Measurements are based on a horizontally arranged capillary filled with the heavy phase in which it is placed a drop of the light phase setting the whole system in rotation. The diameter of the drop which is elongated by centrifugal force correlates with the interfacial tension. Interfacial tension values for the ATPS sample (PEG1000 / MgSO4 / H2O) with several mass proportions are ranging from 0.1 to 1 mN/m. The temperature is kept to 22°C and density (upper and bottom) difference phases are oscillating between 30 to 60 g/cm3. First essays via two couple of porosimetric liquids are based on the previous fundamentals, the two (flux-pressure) curves are obtained when consecutive pores of the membranes, between maximum value dmax (“droplet” point) and minimum value dmin, are opened to flow of the displacing liquid. In this case for the couple isobutanol–water and ATPS (PEG 1000/MgSO4/water) in a specific weight proportion for Whatman-Anopore 0.02 µm inorganic membrane. The resulting porometry presents an S-shaped Fig. 9 aspects ending in a maximum slope (asymptotic permeability), which corresponds clearly shifted to lower pressures for the whole ATPS mixture i.e that the surface tension becomes lower than the typical alcoholic ones. The last recent results concerning to use ATPS into LLDP are being successfully presented for polymeric and inorganic membranes held by IFTS. Especially interesting should be to dedicate more effort to the properties of interesting ternary mixtures, that allowing to analyze smaller pores are also less volatile and not so prone to have important changes in composition along with the experience duration.

Fig. 9: Pressure vs flow curves for ATPS (PEG 1000/MgSO4/water) and isobutanol–water liquid couples [125].

Table 1: Main characteristics and composition of liquid pairs used for LLDP measurements.

| Study | Liquid pair | Composition (v/v) | Interfacial tension (mN/m) | Temperature (°C) |
| --- | --- | --- | --- | --- |
| Grabar and Nikitine [96] | water/isobutanol/methanol | 25/15/7 | 0,35 | 18,7 ÷ 24 |
| Grabar and Nikitine [96] | water/isobutanol/methanol | 5/4/1 | 0,80 | 20 ÷ 23 |
| Peinador et al. [125] | water/isobutanol/isopropanol | 50/35/15 | 0,45 | 20 ÷ 25 |
| Peinador et al. [125] | water/isobutanol/isopropanol | 5/4/1 | 0,9 | 20 ÷ 25 |
| Grabar and Nikitine [96] | water/isobutanol | 1/1 | 2,2 | 18,7 ÷ 21 |
| Calvo et al. [105] | water/isobutanol | 1/1 | 2,0 | 25 |
| Calvo et al. [105] | water/pentanol | 1/1 | 4,80 | 25 |
| Calvo et al. [105] | water/octanol | 1/1 | 8,50 | 25 |
| Phillips and DiLeo [115] | CorrTest© water/PEG 8000/ (NH4)2SO4 | Not defined | ~0,6 | 20 ÷ 25 |
| Peinador et al. [125] | water/PEG 1000/ MgSO4 | 69/19/12 (mass %) | 0,40 | 20 ÷ 25 |
| Peinador et al. [125] | FC-43©/Isobutanol | 1/5 | 4.8 | - ÷ 25 |

To summarize the advantages of LLDP, mostly common to GLDP, we can conclude that this technique: (i) LLDP only accounts for pores open to flux and test membranes in wet state, so closer to real operation way, (ii) GLDP and LLDP cover most of the range of pore usually found in synthetic membranes, being LLDP to analyze pores under the nanometer range [110]; (iii) different immiscible wetting liquid-displacing liquid pairs obtained from the phase demixing of two or three components mixture which can be selected for LLDP measures according to the membrane pore size and swelling affection [97], (iv) LLDP allows also measuring membrane modules in different geometries, including flat, tubular, rectangular or hollow fibers, (v) LLDP technique is an accurate technique able to give the complete PSD information in very reasonable times (~1h for most of the experiences), (vi) the high potentiality of this technique in order to evaluate the active pores in the nanometer and subnanometer range makes convenient to establish a standard procedure. While, on the contrary, some drawbacks should be mentioned: (i) the indirect method, as based on Young-Laplace and Hagen- Poiseuille transport model it must assume parallel tubes pore geometry, (ii) limited accuracy in measuring low flow rates, (iii) correctly selection of the best the porosimetric liquids i.e wetting and displacement ones, (iv) problematic of swelling using wetting alcoholic mixtures, (v) cannot characterize biofouling layer; difficult to correct for adsorbed t-layer as GLDP, (vi) LLDP procedure needs more careful sample preparation and analysis operation, which if not properly done, could easily lead to non-reproducible or inconsistent results.

LLDP has shown accuracy and reproducibility on the characterization of several commercial track-etched polycarbonate membranes in the ultrafiltration (UF) range. In particular a Nuclepore membrane with a nominal pore diameter of 50 nm was analyzed [6]. The quickly performed experiments give a complete picture of the membrane structure in some hours. In this case, the selection of the wetting /displacement liquid permits to obtain accurate porosity and pore densities, using transport models.  Comparison with PSD of the same membrane obtained from image analysis of SEM pictures gives a nice agreement between both results (mean pore radius of 23.6 nm for LLDP results and 24.0 nm for SEM–CIA performed on the same samples used for LLDP). A not so satisfactory agreement is obtained for other indirectly evaluated parameters as porosity (averages of 1.4 and 2.1%) or the mean total number of pores (19.6 × 1012 pores/m2, from LLDP and 11.1 × 1012 pores/m2 for SEM–CIA) due to the uncertainties in the estimation of active layer thickness and the difficulty to obtain a significant number of well-defined pores in a SEM picture.

Similarly, the same precise, accurate, and fast automated LLDP device was used to characterize two series of polymeric flat membranes made from PES and regenerated cellulose [97]. The results showed a nice agreement between different runs, and also very good agreement was found for experiences performed on the same membrane by using different liquid mixtures. Finally, a fair agreement with other authors’ data was obtained. In this case, LLDP results were compared with those coming from Nuclear Magnetic Porosimetry (NMR) which is, in turn, a much more laborious and indirect method that refers to all the pores or voids present within the membranes including those not open to flux. The fair agreement should mean that similar pore size distributions are present within both open and closed pores. Here the PES filters were wetted with isobutanol rich phase while for the NC we found that the membrane sample had been strongly affected by the isobutanol phase, leading to a breaking of the continuous interphase between the active layer and the support, which now appeared as totally detached, it was decided to use ternary mixture Isobutanol/Methanol/water in several proportions for the NC samples.  The possibility to choose among different liquid mixtures, and also to select the role that both liquids play in the porosimetric analysis offers a great advantage and allows analysing many different filters made of different material and so presenting diverse degrees of hydrophilicity/hydrophobicity leading to quite similar results when effective surface tension is assigned for the ternary mixture and to assure a proper wetting of the membrane pores without damaging the membrane structure during the analysis. Another interesting study was performed related to the characterization of virus retention membranes. In that work, it was shown that the maximum pore size obtained by LLDP distributions can be correlated with performance parameters as relevant as bacteriophages log retention value (LRV) or dextran test cut-off values [107, 126]. This correlation clearly fits better when using the distribution’s maximum pore size instead of minimum or mean pore sizes. It allows envisaging possible methods of estimating the retention capabilities of a given filter, without the necessity of performing time-consuming and expensive retention experiments (especially in the case of virus retaining membranes because working with active viruses or bacteriophages requires severe control measurements). The importance of such a straightforward method for the process of new filter development, thus adapted for selected applications, will surely be appreciated by membrane manufacturers and research labs. During the development of appropriate membranes, possible candidates have to be selected. Currently, this is done on the basis of results obtained in a battery of time consuming and expensive solvent retention tests. Instead of those, LLDP could be used, at least for a first approach, to select the most appropriate candidates on which to perform complete characterizations [105]. This can strongly reduce the number of tests and the cost of characterization experiments. Thus, the LLDP method should be recommended for membrane manufacturers and R&D institutions in the first place, moreover, it offers great potential to be developed towards a customer application. Finally, the good correlation between LLDP maximum pore sizes and the other pore size tests suggests the convenience of a simplification of the technique to the easy and rapid detection of maximum pore sizes.

Different porosimetric liquids have been proposed for this kind of filter as other authors have reported swelling of PES membranes when isobutanol/water/methanol mixtures were used [113]  measurements were conducted using an ATPS mixture [115, 126] reducing the swelling into the active layer.  The use of these fluids has several advantages compared to conventional alcohol-water two/three-phase systems typically employed for the pore-size characterization of ultrafiltration membranes. First, both phases are aqueous, so the tendency for membranes to preferentially swell when in contact with one of the fluids is minimized and second, both the PEG and ammonium sulfate are highly soluble in water which greatly facilitates the flush-out of the compounds from the membrane structure after testing. ATPS mixtures also present some drawbacks because they only “work” with hydrophilic membranes being out of order for hydrophobic. Also, during wetting and permeation fluid both phases have viscosities several times higher than water ~ 1 cp, so low UF permeability rate is committed to getting a complete or reproducible PSD for smaller pores.

Retention experiments, in spite of being the default procedure for membrane selection, are not nowadays fully standardized, and then, they can give strongly variable results when performed with different rigs and experimental procedures, in different industrial or research labs. The comparison of pore size distributions obtained from LLDP techniques suggests that PSD strongly correlates with retention capabilities and performance of the membrane, being LLDP a fast and reliable way to estimate MWCO of the analysed membranes. The method used [105] can be standardized to make it an accurate and convenient way to assess possible applications for inorganic or polymeric samples without the necessity to perform complicate and time-consuming retention experiments. This approach is not intended to substitute retention tests as definitive characterization methods for manufacturers deciding the application of their membranes. On the contrary, LLDP can help them to choose better candidates for selected applications among those developed in the corresponding research labs. Another advantage is that the retention test of the real effluents to be separated should be performed on membrane candidates to validate LLDP estimations, but this procedure surely should save many expensive experiments. Table 2 summarizes the LLDP studies for pore size characterization

Table 2: Summary of selected LLDP results.

| Study | Membrane Type and Nominal pores size / MWCO | PSD measured / MWCO | Wetting/ Displacement liquid |
| --- | --- | --- | --- |
| Calvo et al. [6] | PC 0.05µm | 20-80 nm / not tested | water rich phase / İsobutanol rich phase |
| Peinador et al. [97] | PES, RC 5, 10 and 30 kDa | 10-50 nm / 1-100 kDa | Water rich phase/ Methanol-Isobutanol rich phase and Isobutanol rich phase / water rich phase |
| Calvo et al. [104] | PS 20 kDa | 4-20 nm / 1-50 kDa | water rich phase / İsobutanol rich phase |
| Peinador et al. [107] | PES (Virus retention)/ No data | 5-50 nm / 100kDa | İsobutanol rich phase / water rich phase |
| Giglia et al. [126] | PES (Virus retention)/ No data | 5-50 nm / Not defined | Water-PEG 8000 rich phase / water- ammonium sulfate rich phase. |
| Calvo et al. [105] | PES, PS, RC, Zirconium oxide 5-300 kDa | 10-100 nm /1-500 kDa | water rich phase / İsobutanol rich phase |
| Tanis-Kanbur et al. [9] | Aluminum oxide 100 nm, PES 300KDa, PET no data, PTFE no data, PVDF 100kDa and Nylon no data | 80-150 nm for Aluminum oxide, 20-70 nm for PES, 15-40 nm for PET, 60-90 for PTFE, 40-140 nm for PVDF and 50-80 nm for Nylon / 350-1460 kDa | Water-rich phase/ Isobutanol rich phase and vice-versa. |

# Mercury Intrusion Porosimetry (HgP)

Mercury intrusion porosimetry (HgP) is among the characterization methods that are based on the Young-Laplace equation. The main difference with the other two Young-Laplace based techniques (namely, GLDP and LLDP) relies on the fact that, in this case, it is mercury (i.e., a non-wetting liquid) that is forced to enter the pores. The pressure necessary to overcome the interfacial forces can be related to the characteristic radius of each pore by using the previously elucidated Young-Laplace equation (in the form presented in Eq. (3)) and the corresponding contact angle [127, 128]. This equation, when used for HgP, is usually named the Washburn equation [129]. Since contact angles greater than 90º (i.e., non-wetting fluids) lead to negative values of cos θ, the application of Eq. (3) results in negative pressure values. The meaning of this sign simply reflects the fact that mercury, a non-wetting liquid, needs to be forced into the pores.

Experimentally, the setup consists of a pressurized cell containing an adequate amount of the sample to be evaluated. By a stepped increase of applied pressure, mercury is forced to enter all the pores present in the membrane, and the amount of mercury intruded in the system is calculated through measurements of cell conductivity. The plot of intruded volume (more specifically, the differential increment of specific volume) versus applied pressure is usually called a program and contains all information needed to obtain the PSD of the sample. Normally, HgP runs include a complete intrusion-extrusion cycle. In the extrusion part of the cycle, the previously intruded mercury is ejected from the sample during the depressurization step.  All programs, irrespective of their highly variable shape that is dependent on the characteristics of the sample, present two common features, namely, some hysteresis and certain pore entrapment (e.g., dead-ended pores) [9].Usually, only intrusion curve data are used to get PSD, but both curves can be converted into a number of pores. The differential distribution of specific intruded volumes is given by:

(34)

Dvdp=pdpdVdp

where p is the applied pressure, dp the corresponding pore diameter according to the Young-Laplace equation, V corresponds to the specific volume (volume of mercury divided by sample weight) and dV/dp is the derivative of specific intruded volume versus pressure. From this distribution the relative population of the pore in each class can be obtained as follows:

(35)

ndj=Dvdp-dpj-dpj-1dpj-dpj-1/2i=1nDvdp-dpj-dpj-1dpj-dpj-1/22

Of course, for a proper application of Eq. (3), two key parameters, namely, surface tension and contact angle of mercury with the given porous materials, are needed. Values quoted in the literature for mercury contact angle range from 112 to 150º, but most HgP devices include a default value of 130º, which seems to be valid for most of the common porous materials, specifically for most membranes [127]. Regarding air-mercury surface tension, Honold and Skau[130], and some authors after them, used a value of  = 0.474 N/m. Nowadays, a more precise value (i.e.,  = 0.485 N/m) is usually quoted [131], with small variations with temperature.

Many companies market HgP devices since these have broad applications on the characterization of all kinds of porous materials (i.e., not limited to membranes). Some of these companies are Carlo Erba (Macropore), Fisons (Porosimeter P2000), Anton Paar (Poremaster Series), PMI (M/NonM IP), Thermo Fisher (Pascal), Micromeritics (Autopore V) or Quantachrome Co. (PoreMaster). All devices include sophisticated software for automatic control of the measurement and calculation of all porosity-related information, along with appropriate derivation algorithms to get smoothened curves. In most of them, the amount of intruded mercury is determined by the fall of the level of the interface between mercury and compressing liquid. Typically, HgP tests require placing the porous sample into a container (i.e., penetrometer), evacuating the container to remove contaminant vapors, and then forcing the mercury to fill the container. Operationally, the sample is first evacuated and the penetrometer is then filled in the low-pressure port, and then a low-pressure analysis is performed up to the maximum pressure specified, after which the sample is moved to a high-pressure port where a high-pressure analysis is performed up to the higher pressure achievable or selected. Normally, data acquisition software uses appropriate derivation algorithms to get smoothed curves. Maximum pressure of about 60,000 psia (414 MPa) is typical for commercial instruments. This pressure will force mercury into pores down to about 0.003 microns in diameter. The volume of mercury that intrudes into the sample can be determined from measurements of the capacitance. Since the penetrometer is made of glass (i.e., an insulator), filled with mercury (i.e., a conductor), and plated outside with some conductor metal, this combination results in a co-axial capacitor whose capacitance changes as it becomes filled [132]. An interesting feature of this technique lies in the fact that it covers (in a single device) a wide range of pore sizes, usually from 2-3 nm to 1000 microns. Notwithstanding, the smallest pores in this range need intrusion pressures as high as 4500 bar, which for most of the polymeric membranes results in pore structure distortion leading to false results. Due to this reason, HgP is scarcely used for the characterization of polymeric membranes. Another reason for limited use for polymeric membranes such comes from the fact that most such membranes, at least those intended for UF or NF, consist of an asymmetric system formed by a narrow selective layer and placed onto much more open support. When performing an HgP analysis, most of the contribution of mercury intrusion comes from the very porous support, while the volume intruded in the active layer is much smaller such that these pores can hardly be distinguished in the resulting distribution.

Ceramic membranes resist higher pressures and can be easily broken in small pieces that assure the measurement cell present enough number of pores inside to have an accurate analysis, HgP has been quite often used to characterize ceramic materials [133, 134], including ceramic flat-sheet [135] and hollow fiber membranes used in microreactors, membrane contactors or fuel cells [136-140]. Even earlier than GLDP, mercury porosimetry was primarily used for porous characterization of powders and ceramics, later it is found that the method is useful to determine the PSD of membranes. Even though Nakao [5] reviewed several pore size characterization techniques, there was insufficient information about the HgP method. The method was originally proposed by Washburn [127, 129], in a brief note, not including experimental results or any prototype to perform the technique. Using the idea of Washburn, Ritter and Drake in 1945 developed the method, evaluated the contact angle for mercury and designed the experimental procedure which they applied to determine the porosity values of several porous materials like diatomaceous earth, Coors porous plate, Pyrex glass, pelleted gel or activated clay [141]. The first use of mercury intrusion porosimetry to characterize membranes was made by Honold and Skau in 1954, applied to hydrosol type MF membrane filters [130]. After that, many research and industrial applications of the method have proved its reliability on determining pore size distributions, pore shape characteristics, specific surface areas, and porosities. Many of these applications can be found in the excellent review by Rootare [142]. As commented, mercury porosimetry has been used from time ago for characterizing porous materials, especially powders, ceramic materials, and zeolites, materials that can be easily broken in small pieces that assure the measurement cell present enough number of pores inside to have an accurate analysis. HgP, like GLDP, has acquired the recognition of recommended standard for measurements of pore volume and pore size distributions in the macro-mesoporous range (ASTM D4284, D2873, D4404 or DIN 66133) [143]. Most HgP applications that can be found in scientific literature are related with characterization of powders [144, 145], rocks [146], coal [147], zeolites [148], and similar porous materials [131, 149-151]. The newest research on HgP is related to the characterization of porous properties on batteries [152], catalytic materials [153], ceramics for inkjet technology [154], or electrolysers [155]. It has also been used to validate porosity determination procedures [156], or to characterize several types of membranes [157-159]. Finally, it has been used to validate tortuosity calculations for polymeric membranes [160].

# Kelvin Equation-Based Techniques

# Evapoporometry (EP)

In the EP method, the pore size characterization of the membranes is determined via evaporation of a volatile, wetting liquid from the membrane pores by a desktop test setup (Fig. 10a). Some of the authors in this review are among those who have advanced the EP technique [9, 24, 161-164]. The EP setup comprises of an incubator to keep the temperature of the test cell constant (e.g. at 29°C), a microbalance to measure the mass loss from the test cell due to evaporation with time (Fig. 11), a software to record the mass measurements every 10 seconds, and granular activated carbon (GAC) to mitigate the vapour concentration of the volatile wetting liquids outside the test cell. The simplified schematic of the EP test cell, which can be made of aluminium or polytetraﬂuoroethylene (PTFE) to prevent wetting by the volatile liquid, is shown in Fig. 10b. Progressive evaporation of the volatile wetting liquid from the largest pores to the smallest pores can be ensured via supersaturation, which prevails above all the pores smaller than those evaporation-drained ones. Several volatile wetting liquids have been used, such as isopropyl alcohol (IPA), water, 1-butanol, acetone, ethanol, and ethylene glycol [164], with varying performances. In order to maintain the thermodynamic equilibrium saturation at the interface between the gas and liquid phases in the pores during evaporation, EP experiments are conducted in a test cell which has a lid with a small hole (Fig. 10b), whose size can be varied to control the evaporation process [161]. The evolution of the mass and evaporation rate in Fig. 11 represents the data collected during the EP analysis, from which three distinct regions can be demarcated, namely, transient, free-standing liquid evaporation and pore liquid evaporation. The free-standing liquid evaporation rate is constant with time, giving the kinetics of evaporation of the free liquid.

Fig. 10: Schematic of the (a) experimental setup; (b) EP test cell for traditional EP technique, and (c) EP test cell for adaptive EP technique(modified and redrawn from Refs. [9, 24]).

Fig. 11: Representative EP characterization plots of (a) mass versus time and (b) evaporation rate versus time[24].

Both dead-end and continuous pores of the membrane can be characterized by the EP technique. To characterize only the continuous pores, (i) in addition to the volatile wetting liquid, a non-volatile and non-wetting liquid has to be used; and (ii) the test cell is adapted such that both volatile and non-volatile liquids are injected into the test cell via a needle (Fig. 10c) [24].

The method was firstly proposed for the track-etched polycarbonate and aluminium oxide membrane samples in the nominal pore size range of 10-200 nm in 2010 [165]. It was proposed for both flat-sheet and hollow-fiber membranes as a low-cost technique that can also characterize the membrane asymmetry [165, 166]. In addition, the EP technique can provide mass-average and number-average pore size distribution. With respect to these advantages, different membrane pore size characterization studies were performed via the EP method. One of the pioneer studies was conducted by Krantz et al. [34]. An aluminium test cell with a 7.7 cm diameter was designed while the acrylic lid was not used. It must be noted that the acrylic lid is not a mandatory part, but it decelerates the evaporation rate that allows us to better investigate the free-standing layer as shown in Fig. 11. Three different membrane types, which are the polycarbonate (PC), polyvinylidene ﬂuoride (PVDF), and aluminium oxide membranes, were analyzed with different nominal pore sizes vary between 10-100 nm. IPA was preferred as the volatile wetting liquid and the microbalance could handle a mass of 230 g with the nominal resolution and repeatability of 10 and 30μg, respectively. The pore size characterization results were agreed very well with the previously applied indirect and direct methods such as AFM, SEM, FESEM, and LDP, especially for the nominal pore size of 100 nm. Following this, it can be said that the EP can provide accurate pore size characterization without any requirement of transmembrane pressure measurement or a specific and small membrane area for direct observation. The accuracy of the EP technique was also proven for the hollow fiber membranes by comparing the EP and LDP results in Ref. [167]. Polyacrylonitrile (PAN) and PVDF hollow-fibber membranes were investigated by using the IPA and water as volatile wetting liquids. The test cell was made of Teflon® with an inside diameter of 4 cm. Results inferred that both water and IPA presented satisfying results according to the error bars of the replicated tests. For the PVDF membrane, the comparative study deduced that the LDP projected smaller pore diameters than the EP for PVDF due to the membrane compaction and high operating pressures while it projected larger pore diameters than the EP for PAN since the LDP was not able to analyze the pores smaller than 14 nm. In the following study [161], the lid was proposed to decelerate the evaporation rate for better free-standing layer observation and for more accurate results. Aluminum oxide and polyethersulfone (PES) membranes with the nominal pore size of 100 nm were performed via improved test setup design. The lid thickness was 2.7 mm and the hole diameter of the lid varied between 2-10 mm. In addition to different hole diameters for the lid, the impact of different microbalances, which had 1 and 10 μg resolutions with the tare weights of 52 g and 220 g, respectively, were investigated. The test cell was made of Teflon®. It was seen that the microbalance with the higher resolution provided a larger average pore diameter since it was able to resolve the larger pores better. On the other hand, the use of different hole diameters for the lid allowed adjusting the evaporation rate according to the desired conditions. Zamani et al. [163] conducted EP tests for the tubular membranes to characterize the pore size and pore size distribution at the lumen side of the hollow-fibber membranes. A single-bore hollow fibber polyvinylpyrrolidone (PVP)/PES membrane and a multi-bore hollow-fibber PES membrane were used in the experiments. The test cell was made of Teflon® and IPA and 1-butanol were performed as two different volatile wetting liquids. Pores of PVP/PES membrane were smaller than 300 nm while the PES membrane could remove all the particles larger than 20 nm with respect to the membrane manufacturer data. The EP study showed that the results were repeatable and agreed well with one another for both PVP/PES and PES membranes. By means of this study, it was seen that the EP was also able to characterize the lumen side of tubular membranes. In another study, Zamani et al. [164] focused on the investigation of the upper limit of EP-based pore size measurement range. Former studies showed that the lower limit of pore size measurement range was 4 nm and it was directly related to the Kelvin equation. However, the upper limit was related to the vapor pressure depression and relevant evaporation rate; hence, different volatile wetting liquids were performed to see the maximum upper limit for the EP method. Six different volatile wetting liquids, which were water, ethanol, IPA, acetone,1-butanol, and ethylene glycol, were considered at first, but then only IPA and 1-butanol were selected according to the following criteria; (1) as large as possible surface tension-molar volume and (2) high vapour pressure without resulting in very rapid evaporation rate. In addition, it was also mentioned that the liquid should not interact with the deposited foulants or membranes in case of a membrane autopsy whilst the liquid should be non-wetting to the diffusion test cell. The membrane was selected as the aluminium oxide membranes with the nominal pore size values of 100 and 200 nm. The test cell was made of Teflon® and the hole diameter of the lid was 8 mm. It was concluded that 1-butanol presented better pore size distributions at high nominal pore size values thanks to the higher vapour pressure depression that was the cause by higher surface tension-molar volume values. To this end, it was deduced that the upper limit of the EP method can be adjusted by applying different kinds of volatile wetting liquid. Akhondi et al. [162] considered the backwashing process, which is a well-known technique for membrane fouling mitigation, and the impact of backwashing on the pore size distribution was investigated via EP method. PAN and PVDF hollow-fiber membranes had the same membrane properties as Ref.[167]. A microbalance with 10 μg resolution was used and the test cell was made of Teflon®. Deionized water was selected as the volatile wetting liquid. It was inferred that the backwashing achieved better removal performance within larger pores while it was not a strongly effective method for smaller pores. Following the presented developments, Tanis-Kanbur et al. [24] proposed a new and improved technique, which was called as Adaptive Evapoporometry (AEP), to characterize only continuous pores. The traditional EP method characterizes both dead-end and continuous pores; however, the AEP method achieved to characterize only continuous pores by modifying the EP test setup and test procedure. In the AEP method, a non-wetting and non-volatile liquid was used in addition to the volatile-wetting liquid. The volatile wetting liquid was injected into the test setup via the needle from the bottom of the test setup. Once the entire membrane area was wet, the non-wetting and non-volatile liquid were injected with the same needle. During the test, the volatile wetting liquid was evaporated from continuous pores and the pores became empty. By injecting the non-wetting and non-volatile liquid, the membrane pores were plugged up; hereby, the interaction between the empty pores and the pores which were stilled filled by the volatile wetting liquid was blocked. In that study, the volatile wetting liquid was selected as ethanol while the non-wetting and non-volatile liquid were selected as silicon oil since the silicon oil was denser and not mixed with ethanol during the EP test setup. PES and PVDF membranes were used in the experiments with the nominal pore diameters of 28 and 35 nm, respectively. The hole diameter of the lid was 2 mm and the test cell was made of Teflon®. The results stated that the proposed method achieved to characterize continuous pores excluding the dead-end pores and it was seen that the average pore size diameters were found roundly 9.3 and 14.5% less when only continuous pores were taken into account. Besides the EP development studies, a detailed comparative assessments between the EP and LLDP methods were done by using several polymer membranes and aluminium oxide membrane [9]. The main purpose of the study was better understanding the state-of-art for both method and to show the well-agreement between these two techniques even though their theoretical background were based on different theories (see Section 2). In the comparative study, five different polymeric membranes, which were nylon, PET, PES, PVDF, and PTFE, were tested. In addition, an aluminium oxide membrane was investigated in the comparative study. In the EP experiments, IPA was selected as the volatile wetting liquid. The results deduced that the EP presented well agreements for the nylon, PES, and PTFE membranes while there was a difference between two methods for the PVDF and PET membranes. The reason behind that difference was given as the high-pressure requirements of LLDP setup that could negatively affect the membrane structure. Table 3 summarized the EP-based pore size characterization studies.

Table 3: Summary of the EP studies.

| Study | Membrane Type | Nominal pore size range | Volatile wetting liquid |
| --- | --- | --- | --- |
| Krantz et al. [34] | PC, PVDF, Aluminum oxide | 10-100 nm | IPA |
| Akhondi et al. [167] | PAN and PVDF | 35 nm for PVDF, no data for PAN | Water and IPA |
| Akhondi et al. [161] | Aluminium oxide and PES | 100 nm | IPA |
| Zamani et al. [163] | PVP/PES and PES | Pores of PVP/PES smaller than 300 nm, PES membrane can remove the particles larger than 20 nm | IPA and 1-butanol |
| Zamani et al. [164] | Aluminium oxide | 100 nm and 200 nm | IPA and 1-butanol were tested; acetone, ethanol, ethylene glycol, and water were discussed. |
| Akhondi et al. [162] | PAN and PVDF | 35 nm for PVDF, no data for PAN | Deionized water |
| Tanis-Kanbur et al. [24] | PES and PVDF | 28 nm and 35 nm | Ethanol (volatile wetting liquid), silicon oil (non-volatile non-wetting liquid) |
| Tanis-Kanbur et al. [9] | Aluminium oxide, PES, PET, PTFE, PVDF, and nylon | No data for nylon, PTFE, and PET. 300 kDa for PES, 100 kDa for PVDF, and 100 nm for Aluminium oxide | IPA |

# Permporometry (PmP)

Permporometry was firstly proposed by Eyraud in 1984 [168] and modified by Mey-Marom and Katz [169] and Katz [170] in 1986. It has an older history than the EP method in the membrane applications, which have been widely being considered since the early 90s [171, 172]. The experimental setup for the permporometry technique consists of four main sections, namely, gas flow system, diffusion cell, gas chromatograph with data recorder, and temperature control system, as illustrated in Fig. 12.

Fig. 12: Schematic of the permporometry setup (modified and redrawn from Ref. [173]).

The gas flow system consists of condensable gas generators, nitrogen gas, oxygen gas, mass flow controllers, pressure transducers, and gas mixers. For the membrane sample, the condensable gas and nitrogen are flushed at the active layer side, whereas the mixture of a condensable gas, nitrogen, and oxygen is flushed at the supporting layer side. The operating pressure of the diffusion cell can be adjusted to the atmospheric pressure. The condensable gas should be inert, has a reasonable vapour pressure, wettable, and has a high evaporation rate for accurate characterization. As a carrier gas, nitrogen gas is preferred. The reason behind using oxygen is to determine the gas diffusion through the membrane by using gas chromatography. The temperature must be kept constant during the experiments. In the diffusion cell, both membrane sides are flushed with the mixed gases that are saturated with the condensable gas. Thanks to capillary effect (condensation), all membranes pores are filled with the liquid phase of the condensable gas and there is no gas diffusion through the membrane. By decreasing the vapor pressure, gas diffusion proceeds from the larger to smaller pores. At the same time, the oxygen flow is measured via gas chromatography (and data recorder), and thereby the pore size distribution can be plotted. It should be noted that the permporometry technique assumes the pore shape as cylindrical. More details on the permporometry setup and measurement detail can found in one of the cornerstone permporometry studies by Cao et al. [173].

There were different efforts to improve the permporometry method. Schneider and Uchytil [64] proposed an improved permporometry method, the liquid expulsion permporometry, to better characterize the membranes for their smaller sizes and at the high-pressure difference. Moreover, Mourhatch et al. [77] developed a pore network model to take the pore interconnectivity into the account for the permporometry measurements; hence, more accurate pore size characterization was aimed at in the study. On the other, Germic et al. [114] stated that the drying process of the permporometry result may result in a morphological change in membrane structure after comparing the permporometry with different direct and indirect methods such as FESEM and LDP. However, despite drawbacks, the permporometry method has been popularly using in the membrane pore size characterization studies since it characterizes the continuous (active) pores with a simple experimental setup. Cuperus et al. determined the active pore size of an anisotropic ultrafiltration membrane, a polyphenylene oxide (PPO) membrane; and then, the applications were widened for different membrane types such as ceramic, nuclepore, and PSf membranes [171, 174]. Cao et al. [173] also characterized the ceramic membranes with the pore radius range from 2 to 10 nm according to the permporometry and gas adsorption-desorption method; the comparative study deduced that the permporometry technique achieved reliable data for active (continuous) pores while gas adsorption-desorption method was able to characterize both dead-end active pores. Huang et al. [175] performed the permporometry method for the ceramic membranes in a large nominal pore size range, between 2 and 100 nm, by using cyclohexane and carbon tetrachloride as condensable gas. It was seen that the results agreed well with the previous applications and gas adsorption-desorption method. The impact of the condensable gases on the pore size characterization was performed by Tsuru et al. [176]. Water, methanol, ethanol, carbon tetrachloride, hexane, and isopropanol were compared to one another in the permporometry technique. The comparative study inferred that the water, hexane, and carbon tetrachloride showed similar pore size distribution trends for the average pore size values above 1 nm. Another study of Tsuru et al. [177] focused on the comparative pore size distribution trends of helium and nitrogen as gas carriers while the condensable gas was selected as water in the ceramic membranes in the average pore size range between 0.5 and 2 nm. For large pore size values (as large as 2 nm), the pore size distribution trends of helium and nitrogen were seen in a good agreement with each other whilst the pore size distribution trends of helium were shifted to smaller pores compared to the pore size distribution trends of nitrogen for the average pore size values. Hence, it was stated that the micro pores allowed only helium permeation. Besides the aforementioned works, there have been a lot of membrane fabrication studies that preferred the permporometry for the pore size characterization. Table 4 summarizes the remarkable permporometry-based pore size characterization studies.

Table 4: Summary of the remarkable permporometry studies.

| Study | Membrane Type | Purpose of Permporometry |
| --- | --- | --- |
| Cuperus et al. [171] | Anisotropic polyphenylene oxide (PPO) membrane | Pore size distribution, pore size, and top layer thickness were explained via theory and experiments. |
| Cuperus et al. [174] | PPO, Nuclepore, ceramic, and PSf membranes | Comparative permporometry study on various membrane types to better determine the limitations of the method. |
| Cao et al. [173] | Ceramic membrane | Performance assessment in the pore size range between 2 and 10 nm. |
| Huang et al. [175] | Ceramic membrane. | Performance assessment in the pore size range between 2 and 100 nm. Comparative work with the gas adsorption-desorption method. |
| Germic et al. [114] | Polyacrylonitrile membrane | Comparison of permporometry with different methods: SEM, LDP, gas adsorption-desorption. |
| Kim et al. [25] | Regenerated cellulose and polysulfone | Comparative thermoporometry and permporometry study were conducted in the nominal pore size range of 2.6 and 14.1 nm. |
| Tsuru et al. [176] | Lab-based ceramic membranes (sol-gel membranes) | Various condensable gases were carried out: water, methanol, ethanol, isopropanol, carbon tetrachloride, and hexane. |
| Tsuru et al. [177] | Lab-based ceramic membranes | Pore size distributions of nitrogen and helium-based permporometry studies were compared. Water was selected as the condensable liquid. |
| Nishiyama et al. [178] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic membrane. Helium and nitrogen were the condensable and non-condensable gases, respectively. |
| Chowdhury et al. [179] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic membrane. Hexane was used as the condensable gas. |
| Wang et al. [180] | Lab-based ceramic membrane | The defect size distribution was characterized for different ceramic membrane types. |
| Hedlund et al. [181] | Lab-based ceramic membrane | Permporometry was used to characterize the ceramic membrane types with a different film thickness range from 300 to 1800 nm. |
| Higgins et al. [182] | Lab-based ceramic membrane | Helium and nitrogen were used as the condensable and non-condensable gases, respectively. |
| Chowdhury et al. [183] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic (zeolite) membrane. |
| Yang and Li [184] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic (tubular TiO2/Al2O3) membrane |
| Kuhn et al. [185] | Lab-based ceramic membrane | Characterization of the lab-based manufactured multichannel ceramic (zeolite) membrane |
| Schillo et al. [186] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic (alumina) membrane during the rapid thermal processing |
| Lindmark et al. [187] | Lab-based ceramic membrane | Permporometry was used to investigate the pore size distribution during the selectivity enhancement process of the zeolite membranes |
| Hofs et al. [188] | Commercial ceramic and PES/PVP membranes | Characterization of the ceramic and polymeric membranes for permeability and fouling assessments. |
| Zeidler et al. [189] | Lab-based ceramic membrane | Characterization of the low molecular weight cut-off ceramic membranes |
| Karimi et al. [190] | Lab-based ceramic membrane | Characterization of the zeolite membranes during the defect blocking process |
| Blumenschein et al. [191] | Commercial ceramic membranes | Characterization of the ceramic membranes to better describe the rejection curves during the organic solvent nanofiltration |
| Chang et al. [192] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic (zeolite) membrane during the rapid thermal processing |
| Simon et al. [193] | Lab-based ceramic membrane | Characterization of the lab-based manufactured ceramic (zeolite) membrane during the micro-defect sealing by chemical vapor deposition |

# Gas Adsorption-Desorption (GAD)

Among the techniques of membrane characterization based on the Kelvin equation, the Gas Adsorption-Desorption (GAD) technique is one of the most commonly used, which after appropriate modelling is able to supply information about the PSD of microporous and mesoporous membranes, along with other types of porous materials. The technique is also known as nitrogen physisorption [194] and low-pressure nitrogen gas adsorption (LP-N2GA) [195]. Most of the GAD studies deal with porous materials (e.g.,  coals, carbons, other rocks) other than membranes [196-198]. However, compared to the permporometry, the use of the gas adsorption-desorption method is less since the permporometry method is fundamentally developed by the same theoretical background and characterizes only continuous (active) pores while the gas adsorption-desorption method characterizes both dead-end and continuous pores.

The GAD analysis is done with the well-established and commercial gas sorption analyzers. A test cell, which holds a membrane sample, is located in the analyzer, and the pore size characterization is done by determining the gas volume that is sent to the sample. Before the experiments, the membranes are dried at high temperatures (i.e., values that do not affect the membrane negatively). The selected gas must be non-condensable at the liquid nitrogen temperature. The gas flow, the gas of which is generally helium, is sent to the sample at a known flux value, and the pressure measurements are done concurrently to determine the dead volume of the membrane sample. Then, the sample discharging process is done to reach the maximum vacuum value, and the constant-flux nitrogen flow is sent to the sample and reference cell. whilst the pressure and time data are collected in parallel to the experiments. The comparison between the sample and reference cells allows us to analyse the adsorption isotherm of the sample. More details on the gas adsorption-desorption setup can be found in earlier reports [194, 199].The GAD technique is used to determine the surface area of porous materials, based on the determination of the BET isotherm. The resulting method gives the so-called BET surface area. It is common to use nitrogen as the adsorbate, though other gases could be used depending on the membrane adsorption characteristics [200]. In the following, the GAD procedure is presented with the required assumptions and calculations.

# Determination of Surface Area

The use of the Brunauer-Emmett-Teller (BET) method is generally considered the standard to obtain the whole surface area of a porous sample. The method, based on the adsorption kinetics model [201], relies on several assumptions whose validity must be ensured: (1) in all layers except the first, the molar enthalpy of adsorption corresponds to the enthalpy of condensation, L; (2) in all layers except the first, condensation-evaporation conditions are equal; (3) at the saturation pressure (i.e., pr = 1), all gases are condensed on the solid surface tending to an infinite number of adsorption layers. Based on the validity of such assumptions, Brunauer et al. [201] obtained a simple linear relationship between adsorbed volume and pressure:

pVp0-p=1VmC+C-1VmCpp0		     (36)

where Vm is the total volume adsorbed in the monolayer per unit mass, V the volume adsorbed per unit mass, and C a constant related to the isotherm shape. According to Eq. (6), a plot of p/V(p0-p) versus pr results in a straight line, whose fitting allows us to obtain the two unknown parameters, namely, C and Vm from the slope and the intercept, respectively. To ensure the validity of this model, which works nicely for relatively low pressures, only results corresponding to relative pressures in the range of 0.05 to 0.3 can be fitted. The final calculation of the BET surface area is related to a geometrical consideration for the adsorbate molecule. Usually, this molecule is nitrogen, and thus using a value of 16.2 Å2 for molecular cross-sectional area Am of the nitrogen molecule, the surface area can be calculated by:

A(BET)=VmVgNAAm		     (37)

where Vg is the gas molar volume at STP conditions and NA the Avogadro number.

# Determination of Mean Pore Size and Pore Size Distribution

In contrast to the BET surface area determination, whose procedure is well known and used in many cases in the same manner (as a recognized standard for surface area calculation), the use of GAD to determine the porous characteristics of a certain material lacks a uniform approach and is highly dependent on the shape of the adsorption isotherm, which is strongly tied to the presence of mesopores (i.e., between 2 and 50 nm in diameter) and micropores (i.e., lower than 2 nm in diameter). In the case of macropores (i.e., larger than 50 nm in diameter), the use of GAD leads to operation at pressures very close to the saturation one (pr = 1), which makes inaccurate any calculation of such pores.

- Mesopores Analysis:

In the case of mesopores, the adsorption of gas onto porous materials is mostly due to capillary condensation. The adsorbate molecules (except the first adsorbed layer) can be considered as a liquid in such a way that the interphase between gas and adsorbed molecules is at vapour-liquid equilibrium. As it is well acknowledged, such equilibria are determined by the meniscus curvature, and this fact can be related to the geometry of the pore via the Kelvin equation, which can be expressed as:

rk=2γVlRTln(pr)		     (38)

where rk is the radius of the equivalent hemispherical meniscus (or Kelvin radius), γ the surface tension, and V1 the molar volume of the liquid condensate.

Supposing cylindrical pores, the pore radius is given by rp = rk+ t, where t is the correction made to rk to account for the thickness of a layer already adsorbed on the pore walls. The accuracy of the mesopores analysis strongly depends on the calculation of t. Ideally t should be determined from a measurement of adsorption onto a completely flat surface from the same material under study; however, such measurement is not practical (and at times impossible) for most cases. Therefore, some sort of phenomenological correlation should be used. Among the correlations developed to address this problem, the Halsey correlation [202, 203], obtained from abundant measurements on different flat materials, seems to give goods results for most samples with mesopores. This correlation gives t, in Angstroms, as:

t=0.3545lnp0p1/3		     (39)

The same correlation can be used also for pores having a parallel-sided slit configuration, in which case the slit-width, dp is given by dp = rk + 2t. By using this correlation, a number of computational procedures have been proposed to deal with mesopores from adsorption experiments. One of the most known is that proposed by Dollimore and Heal for cylindrical pores [92, 204].

Finally, it should be noted that both isotherms, during adsorption and desorption, can be used to determine the PSD of a mesoporous membrane, but they do not match exactly due to certain hysteresis (e.g., attributed to the presence of pore interconnectivity [205]) in the filling process. Normally, the adsorption isotherm is considered more precise to determine the PSD of the sample.

- Micropores Analysis:

Unfortunately, the Kelvin equation loses validity for micropores, because substitution of the capillary condensation mechanism by multilayer adsorption leads to a distortion of the usual mesopores isotherm shape. Approaches proposed to address this range from simple reformulations of the Kelvin equation to more complex descriptions of the process. Regarding the reformulation of the Kelvin equation to account for the adsorbed layer thickness, it has been shown to give inconsistent results[206]. Methods like Dubinin-Radushkevich [207-209] or Horväth-Kawazoe [210, 211]  have been used to obtain information about micropores. One of the most interesting ones is the HgP-method, which is based on the t-plot developed by Lippens and De Boer[212]. These methods depart from a plot of the adsorption isotherms in terms of the adsorbed layer. Here, t values can be obtained by using an empirical correlation by Harkins and Jura [213]:

t=13.990.034+lnp0p1/2		     (40)

# Determination of Overall Porosity

Apart from determining the BET surface area, and getting the PSD of a porous sample, GAD can be also used to determine the porosity of the membrane. For such purpose, a quite simple procedure has been developed by some of the authors of this review [214, 215]. The method is based on determinations of the pressure of Helium gas introduced into the sample holder as a function of time. Working at low enough pressures (recommended values range between 7 and 16 kPa), helium can be considered to behave as a perfect gas. So, the following equation can be used to fit the experimental data:

p=k1+k2Vg		     (41)

where k1 and k2 are unknown constants to be determined from experimental data-fitting and Vg the volume occupied by the gas. For practical purposes, the measurement consists of the determination at three different situations (namely, empty holder, a non-porous sample of known volume, and the membrane sample), from which the actual solid volume of the sample can be determined. From the geometric volume of the sample (accounting for the maximum dimensions and not the existence of inner pores), the ratio of both volumes gives the sample porosity.

The reason for using helium as the intruding gas is its small molecular size and the fact it does not adsorb onto the holder walls and on the pore walls at the operating temperature (i.e., 77 K, since the holder is placed in a Dewar containing liquid Nitrogen). At the same time, nitrogen gas is introduced in the reference holder. The saturation pressure measured in this holder serves to normalize the sample pressure, thus avoiding the influence of atmospheric pressure variations.

The market offers a range of technical solutions and instruments for gas adsorption analysis, including the 2020 and 2020 Plus Series of Micromeritics, the wide range covered by Autosorb, NOVAtouch, and Quadrasorb Series developed by Quantachrome, the Belsorp series (from BEL Japan), 3P Inst. (formerly Quantachrome GmbH developed the 3P micro 300 and 3P meso 400, aimed for micro and mesopores analysis, respectively), and the Omnisorp Series by Coulter (not continued later by Beckmann-Coulter). Apart from design and robustness, the main differences between these types of equipment come from the method used to determine the amount of gas adsorbed onto the sample, which can be categorized as three main groups: (a) volumetric, which measures the volume and pressure of the gas at equilibrium conditions (both static and continuously); (b) chromatographic; and (c) gravimetric, which determines the variation of the solid mass with the aid of an accurate microbalance [4].

GAD method is generally used in comparison studies which consist of both indirect and direct techniques. These studies devoted to the characterization of several kinds of membranes [216-220] and supports [221]. In the inorganic membrane applications, Cini et al. [222] applied the gas adsorption-desorption technique for the low surface area and macroporous ceramic membranes to characterize pore size and understand the fouling phenomena. Bailey et al. [223] performed a comparative pore size characterization study for the ceramic membranes by using the gas adsorption-desorption and atomic force microscopy which represented the indirect and direct techniques. Basumatary et al. [224] developed a ceramic membrane for the separation of chromic acid from aqueous solutions, and they characterized the pore size of the developed membrane by using gas adsorption-desorption method, XRD, and FESEM techniques. Maghsodi et al. [225] proposed a hydrothermal ultrasonic method to remove arsenic from the aqueous solution via a ceramic membrane type which was characterized by using the gas adsorption-desorption method. Apart from the inorganic membranes, the gas adsorption-desorption technique has also been using for the characterization of organic membranes. Prádanos et al. [199] focused on the analysis of the internal surface area, volume, surface, and the number of pores for the microporous symmetric membranes by combining the atomic force microscopy and the gas adsorption-desorption technique. Following that study, two polyethersulphonic microporous composite membranes with two different nominal molecular weight cut-off values of 4000 and 30000 Da were characterized by comparing the gas adsorption-desorption, LDP, and atomic force microscopy. Hereby, the pore size characterizations via Young-Laplace equation-based LDP, Kelvin equation-based gas adsorption-desorption, and direct observations were compared to one another [216]. The same group also compared the surface and bulk porosities of ten different polycarbonate membranes with respect to the indirect (gas adsorption-desorption) and direct (atomic force microscopy) techniques [219]. For gas separation purposes, Zhang et al. [226] developed a nanoporous carbon membrane and the characterization was carried out by observing the indirect (gas adsorption-desorption) and direct (FTIR, TEM, SEM) methods. A more theoretical approach can be found in the recent work by Zeng et al. [227], in which authors study by a mathematical simulation the microscopic mechanisms of adsorption and desorption in uniform-sized pores, with the aim to elucidate the origin of hysteresis, especially in dead-ended pores.

# Gibbs-Thomson Equation-Based Techniques

# Thermoporometry (ThP)

The Gibbs-Thomson-based measurements are thermoporometry measurements that characterize the pore size and pore size distribution according to the phase-change curves (thermograms) via calorimeters. Two thermograms, namely, solidification and fusion, are used in the thermoporometry tests. Before the experiments, the sample is frozen and melted in sequence to obtain a crystallized phase. This process also prevents thermal flash because the initial overcooling propagates the penetration of the solid into the pores [228]. Like the gas adsorption-desorption technique, thermoporometry is also one of the oldest pore size characterization techniques for porous mediums [35]. The methodology and theory were firstly proposed and validated for powders that were porous mediums; then, it was better described for the membrane pore size characterization studies [228]. After presenting the thermoporometry technique, one of the pioneer studies was done by Zeman et al. [229] by characterizing the polysulfone ultrafiltration membranes. Following that study, Quinson et al. [230, 231] characterized the polycarbonate membrane with the average pore size diameter 15 nm to focus on the swelling phenomena. Broek et al. [232] conducted thermoporometry for the hollow fibber hemodialysis membrane that was based on cellulose to characterize the pore structure in the artificial kidney studies. Cuperus et al. [233] focused on the thermoporometry tests for both ceramic (alumina) and polymeric (polyphenylene, polysulfone, and cellulose acetate) membranes to see the feasibility of the method on different membrane types that had different casting thicknesses. It was seen that the method was reliable for isotropic membranes while there was a doubt for the anisotropic membranes since the pore in the sublayer of anisotropic membranes may prevent accurate characterization. Similar to Ref. [225], Terrazas-Bandala et al. [234] also aimed to remove arsenic from aqueous solutions. Cellulose acetate and active carbon membranes popularly for that purpose and their characterization procedures were completed via the thermoporometry method. All the aforementioned studies are also summarised in Table 5. It can be said that whole thermoporometry studies were conducted by using water although the method can also be applied with different kinds of liquids. The main reason behind this fact is the easy observation of water-ice phase change plots via thermograms.

Table 5: Summary of remarkable thermoporometry studies in the literature.

| Study | Membrane type | Liquid medium |
| --- | --- | --- |
| Brun et al. [35] | Powders 1 | Water and Benzene |
| Zeman et al. [229] | Polysulfone | Water |
| Quinson et al. [230] | Polycarbonate | Water |
| Quinson et al. [231] | Polycarbonate | Water |
| Broek et al. [232] | Cellulose | Water |
| Cuperus et al. [233] | Ceramic (alumina), and polymer (polyphenylene, polysulfone, and cellulose acetate) | Water |
| Jallut et al. [228] | Polysulfone | Water |
| Terrazas-Bandala et al. [234] | Active carbon and acetate cellulose | Water |

1: They are not membrane samples, but the study is listed since it was the first study of the detailed theory.

# Spectroscopy-Based Techniques

In addition to the equation-based indirect methods, there are also spectroscopy-based indirect techniques. Compared to the equation-based methods, they are relatively new and they are based on the images that are obtained from the spectroscopy measurements. In fact, the spectroscopy can be considered as a direct technique, but the collected images can be used in mathematical expressions to estimate the pore size and characterize the membrane. The equation-based indirect methods are useful and accurate approaches fort he MF and UF membranes but they are not able to estimate the pore size for tighter membrane types such as NF and RO membranes. The spectroscopy-based techniques can collect high-resolution images that allow measuring the membrane pores. Then, the measured pore data is used in indirect calculations. Below, we reviewed two different spectroscopy-based methods which are the positron annihilation lifetime spectroscopy (PALS) and the synchrotron radiation (SR). Even though there are few studies about them on the membrane characterization studies, their accurate measurements show that they are promising techniques for sub-nanometer scale membranes. Thus, the below-mentioned parts can be useful for the readership.

# Positron Annihilation Lifetime Spectroscopy (PALS)

Up to now, the PALS method has been exclusively applied to the low-pressure & seawater RO and tight NF membranes. It is worth mentioning that, in these kind of studies, holes or voids are more used words, rather than pores, because they surely describe more adequately the very small interstices between polymer molecules which play the role of pores in so tight membranes. Fujioka et al. [36] summarized the main characteristics of the tested RO and NF membranes in a comprehensive review study. In the polymeric membrane characterization studies, the beam intensity was applied in the range of 1.0-2.0 keV that corresponds to the mean positron implementation depth range of 40 to 200 nm with the material density of 1 g/cm3 [36]. Kim et al. [235], applied the PALS to a low-pressure polyamide RO membrane and determined the pore radius as 0.217 nm. Tung and his colleagues [20, 23] performed the PALS to tighten NF membranes. In Ref. [23], a polyamide NF membrane was characterized with the beam intensity of 2.0 keV, and the pore radius was calculated as 0.279 nm. In the following study [20], the same membrane type was fouled first; then, the PALS was applied with the same bean intensity (2.0 keV) and the pore radius was determined as 0.312 nm. Chen et al. [236] investigated the correlation between the free-volume holes and the molecular separation in the RO process and conducted the PALS analysis for a low-pressure RO and two NF membranes with the beam intensity of 2.0 keV. For the low-pressure RO membrane, the hole radius was calculated as 0.203 nm whilst it was 0.265 and 0.275 nm for two different NF membranes, respectively. Fujioka et al. [237] evaluated the rejection of uncharged chemicals during the RO process and investigated three different polyamide RO membranes via the PALS method. Two of the membranes were low-pressure RO membranes and the other one was the seawater RO membrane. The beam density was 1.0 keV for all experiments. The PALS study inferred that the seawater RO membrane had a smaller hole diameter with 0.259 nm whereas the low-pressure RO membranes had a hole diameter of 0.289 nm; hence, the seawater RO membrane showed better rejection performance than the low-pressure RO membranes. Another PALS study by Fujioka et al. [238] analyzed the impact of heat treatment on the membrane pore size characterization for RO membranes. It is a known fact that heat treatment improves the rejection and reduces the fouling, but it also results in a decrement of water permeability. Therefore, the impact of heat treatment on the resulting pore size characterization was a critical question. The PALS study was carried out with a beam intensity of 1.0 keV for two different low-pressure RO membranes. The hole radius was calculated by 0.267 nm for both membrane types, and it was seen that the heat treatment has a negligible impact on the pore characterization. Apart from the heat treatment effect, the impact of humidity on the pore filling was analyzed by Lee et al. [239]. A polyamide seawater desalination RO membrane was used in the experiments and the hole radius was found as 0.240 nm. It was deduced that the pore swelling mechanism was observed by increasing of humidity while the other operating parameters were maintained constant. The recent studies on the PALS analysis has focused on different membrane types. Fan et al. [240] characterized the polyvinyl alcohol-graphene oxide nanocomposite membranes with different weight ratios of grapheme oxide and polyvinyl alcohol. The weighting ratio was in the volume range of 0.15-0.72 %. The results showed that the hole radius varied between 0.219-0.222 nm despite a wide weighting ratio range for the nanocomposite membranes. Differently than the RO and NF membranes, Kim et al. [241] developed cellulose triacetate and polyamide forward osmosis (FO) membranes and compared the pore size with the seawater RO membrane. Pore size diameter was determined with the PALS method whereas the pore size distribution was obtained via the fractional rejection method. The comparative PALS studies between the FO and seawater RO stated that the FO had a higher hole diameter value with the diameter value of ~0.29-0.30 nm than the seawater RO membrane with the hole diameter value of 0.20-0.24 nm. Besides the aforementioned studies, the PALS has also carried out for the inorganic membranes. Ma et al. [242] prepared four zeolite membrane types with different microstructures on alumina supports in a lab environment and conducted the characterization according to the PALS technique. The hole radius values were varied between 2.73 and 3.64 Å.

# Synchrotron Radiation (SR)

SR technique has some advantageous characteristics such as wide range of the spectrum, high intensity, high polarization, small source so that the high brilliance; it can be a beneficial approach for the applications in the fields of materials characterization and analysis, chemistry, micromechanics, biology, physics, geology, archaeology, and medicine [243, 244]. SR-based X-ray diffraction (XRD) and small-angle X-ray scattering (SAXS) characterization methods provide performance-related information about morphology and structural differences under operating conditions [245]. Following these advantages, the SR method has been applied to polymeric and inorganic membrane characterization studies to test the membrane performance since the beginning of the 2000s. Lassinantti Gualtieri et al. [246] used the SR technique as a part of high-temperature X-ray powder diffraction (HT-XRPD) study for testing both zeolite film and alumina support of the MFI zeolite membranes to understand the crack formation during the calcination process. The porous alumina support was found expanded, and the MFI zeolite crystals were contracted during calcination heating. Jeong et al. [247] studied the synchrotron X-ray radiation technique to investigate the microstructure of siliceous ZSM-5 (MFI) zeolite membrane during the calcination process. After the calcination process, the MFI crystals of the membrane were found under compressive stress and their thermal behavior is different from those in powder. This compressive stress might be one of the reasons that crack forms in the membrane which affects overall membrane performance. Barranco-Garcia et al. [248] investigated the gas transport phenomenon of isotactic polypropylene (iPP) based nanoporous gas separation membrane with the changing amounts of SBA-15 silica (iPP/SBA-15) via SAXS SR technique. Confinement effect was monitored, and the silica particles played a filler role which enhanced the rigidity and lowered the deforms. Other studies were performed to investigate the depth-dependent structure during permeation performance and the effect of ethanol swelling of poly(urea-urethane) (PU) membranes via small-angle X-ray scattering (SAXS) SR technique [249, 250].

Apart from those studies, Remigy and Meireles [251], used the SR computed microtomography method to study the pore characterization and 3-D architecture of polysulfone (PSf) and poly(vinylidene fluoride-co-hexafluoropropylene) (PVDF-HFP) hollow fiber microfiltration membranes. Pore geometry and morphology of the membranes were observed by analyzing the 2-D images and reconstruction of 3-D architecture via commercial visualization software. Lee et al. [252], studied the structure of nanoporous polymeric membranes via synchrotron X-ray tomography imaging. Membrane porosity, pore channel and interconnectivity, 3D structure, and the tortuosity characterizations were achieved by performing 3D image analysis. The results inferred that nanoscale imaging can be reliably preferred in the membrane morphology-based mass transfer investigations.

The above-mentioned studies show the possibility of the synchrotron radiation applications for the membrane characterization works, but the number of pore size characterization based studies via SR technique are relatively low compared to other methods. The pore size characterization of the MF and UF membranes have already had well-established techniques, but more detailed and deep-understanding techniques are required for the nanometer-scaled membrane types. Therefore, synchrotron radiation-based studies can be carried for this type of membranes for their accurate pore size estimation and characterization.

# Conclusions and Future Directions

The present study reviewed the existing indirect porosimetric membrane characterization techniques for the analysis of PSD in polymeric and ceramic membranes. These techniques have been thoroughly analysed with respect to their theoretical backgrounds, measurement details, experimental procedures, and remarkable applications, as can be found in the scientific literature. The methods were classified as Young-Laplace equation-based, Kelvin equation-based, and Gibbs-Thomson equation-based, and spectroscopy-based techniques. The bubble point test, gas-liquid displacement porosimetry, liquid-liquid displacement porosimetry, and mercury intrusion porosimetry were reviewed under the epigraph of Young-Laplace equation-based methods whereas evapoporometry, permporometry, and gas adsorption-desorption were reviewed as Kelvin equation-based methods. Thermoporometry technique was explained under Gibbs-Thomson equation-based methods. In addition, positron annihilation spectroscopy and synchrotron radiation were examined jointly as spectroscopy-based techniques. The outcomes of the present review were as follows:

- Liquid Displacement Porosimetry is a simple method that provides a flow-based pore size and PSD for larger membrane areas with minimum sample preparation necessity and small lab footprint. On the other hand, it is limited to MF (gas-liquid interface) and UF (liquid-liquid interface) membranes, it requires high pressure for the smaller pores (NF range) detection, and assumption of pore geometry is needed for the complete pore number calculation. Nevertheless, an accurate and very informative flow-based PSD can be obtained without any geometry assumption. Mercury Intrusion Porosimetry is a volume-based pore size and PSD technique and it can be used for the characterization of many porous materials such as powders, rocks, ceramic membranes, catalytic materials with minimum sample preparation. However, the requirement of high pressure for the detection of smaller pores makes it barely used technique for the polymeric membranes.

- Evapoporometry offers a mass-based pore size and PSD measurement at the membrane surface, it can be used for larger membrane sizes. Tests can be conducted at ambient conditions with minimum sample preparation and small lab footprint. The method can characterize the continuous pores and pore connectivity in polymeric membranes. It can also detect the effect of fouling. However, it is limited to MF and UF membranes and it creates a problem for the samples with low porosity.

- Permporometry gives a flow-based pore size and PSD results at the membrane surface. Larger membrane areas can be tested with minimum sample preparation. It can also measure the continuous pores. On the other side, it is limited to MF and UF membranes, it requires pore geometry assumption, needs partial pressure and flow rate measurement, and troublesome for the samples with low porosity.

- Gas Adsorption-Desorption presents a volume-based pore size and PSD at the membrane surface for larger membrane samples, and it can be used for MF, UF, and NF membranes with minimum sample preparation. The disadvantages of the method are pore geometry assumption, expensive equipment, restricted accuracy of pressure, and volume measurement.

- Thermoporometry measures a mass-based pore size and PSD for larger membrane sizes with minimum sample preparation and lab footprint. The main drawbacks of the method include pore geometry assumption, measurement of the heat, and it is troublesome for the samples with low porosity.

- Positron Annihilation Lifetime Spectroscopy is a spectroscopy-based indirect method that is an especially useful technique for tight NF and RO membranes. It measures the pore size by using the measurements from high-resolution images. The main limitations are related to the membrane test sample which must be vacuumed and form positronium formation.

- Synchrotron radiation obtains high-resolution radiographs that are used to analyze membrane pore characteristics via two-dimensional or three-dimensional image analysis. The time-consuming image analysis and raw data assessment are the main drawbacks of the SR method.

Any of the reviewed methods can be useful to have complete information about the pore size and porosity of our membranes. In that sense, LLDP and Evapoporometry are considered the most complete and precise methods giving an accurate view of actual PSD for any membrane in the MF-UF range, while the extension to NF and RO membranes can be done via spectroscopy-based techniques (e.g. PALS) or some other direct techniques as microscopic ones. And moreover, those porosimetric methods, if being very useful for preliminary selection of samples for selected applications, such selection must be complemented and tested through appropriately conducted test retention experiments.

One major drawback of all those PSD characterization techniques is related with actual microstructure of a porous medium which in many cases is quite different from a simple bundle of parallel cylindrical tubes. Porous systems are rather characterized by a system of more or less interconnected pores. Several studies dealing with the pore network simulation (PNS) method of fluid displacement [253-257] are inspired on the classical invasion-percolation algorithm. It combines the tracking of the fluid-fluid interfaces and the invasion of the throats (corresponding to the constrictions of the pore space) whose threshold capillary pressures, as given by the Young-Laplace equation are lower than the existing pressure difference between the two fluids [257]. In the case of PNS simulation of EP it is made by combining the computation of vapor diffusion in the fraction of the pore space invaded by the gas phase presenting in the network and the identification of the throat of lower capillary pressure invasion threshold [257].

As a result, the information on the PSD obtained from LLDP and EP techniques can be questioned in some cases and must be assessed for improved accuracy by comparison between experimental results and PNS calculations.

.

Although the indirect methods are established techniques for different membrane types, there are still research gaps for their further improvements which can be mentioned. For the case of MF and UF membranes: all Young-Laplace, Kelvin, and Gibbs-Thomson-based techniques can be improved via the use of different liquids. This could be useful to assess the impact of liquid thermophysical properties (e.g. surface tension, wettability) on the pore size estimation since interaction between liquid and membrane material can slightly or dramatically change the pore size distribution plots. Furthermore, the performance measurements of these techniques are subjective, which means the average pore size distribution and pore size estimation are determined by engineers or operators. To make an objective assessment, the machine learning algorithms can be developed for data management and plot fittings of the pore size distributions. In this way, membrane characterization can be predicted with a faster and reliable way. For the nanometer and sub-nanometer sized membranes, PALS and SR techniques have been gaining importance since they can provide a deep understanding of the hole structures. Even though the PALS studies provide reliable pore size characterization for the sub-nanometer-scale membrane structures, the relation between the fouling mechanism and pore size characterization can be carried out in detail with other membrane shapes- and volume-based properties. Also, the mechanistic understanding of the membrane fouling can achieve a more detailed assessment of this kind of membrane types. For this purpose, the molecular dynamics simulation studies can be performed to see the relationship between the transport properties and pore size characterization. Similar to the PALS technique, the molecular dynamics simulation can also provide a deeper understanding of the SR measurements.

# Acknowledgments

M. B. Tanis-Kanbur and J.W. Chew thank the Singapore Ministry of Education Tier 1 funding (2019-T1-002-065). J.I. Calvo and A. Hernández want to thank the Regional Government of Castilla y León (Projects: CLU2017-09, UIC082 and VA088G19) and the Spanish “Ministerio de Ciencia e Innovación (MCINN)” (project: MAT2016-76413-C2-1-R) through EU-FEDER funds. R. I. Peinador thanks the French “Ministère de l’Enseignement supérieur, de la Recherche et de l’Innovation“  funding (Crédit Impôt de la Recherche CIR-IFTS-2019).

# References

[1]	G. Owen, M. Bandi, J. A. Howell, and S. J. Churchouse, "Economic assessment of membrane processes for water and waste water treatment," Journal of Membrane Science, vol. 102, pp. 77-91, 1995/06/15/ 1995, doi: https://doi.org/10.1016/0376-7388(94)00261-V.

[2]	 S. Loeb, "LOEB-SOURIRAJAN MEMBRANE: HOW IT CAME ABOUT," in ACS Symposium Series, 1981, pp. 1-9. [Online]. Available: https://www.scopus.com/inward/record.uri?eid=2-s2.0-0019662223&partnerID=40&md5=1124c041240e9e2a2480219d264683bf. [Online]. Available: https://www.scopus.com/inward/record.uri?eid=2-s2.0-0019662223&partnerID=40&md5=1124c041240e9e2a2480219d264683bf

[3]	R. Bai and H. F. Leow, "mufiltration of polydispersed suspension by a membrane screen/hollow-fiber composite module," Desalination, vol. 140, no. 3, pp. 277-287, 2001/11/20/ 2001, doi: https://doi.org/10.1016/S0011-9164(01)00377-0.

[4]	A. B. J.I. Calvo, P. Prádanos, L. Palacio  and A. Hernández, "Membrane Characterization: Porosity," in Encyclopedia of Membrane Science and Technology. New York (USA): Wiley Interscience Pub, 2013, pp. 1-35.

[5]	S.-i. Nakao, "Determination of pore size and pore size distribution: 3. Filtration membranes," Journal of Membrane Science, vol. 96, no. 1, pp. 131-165, 1994/11/28/ 1994, doi: https://doi.org/10.1016/0376-7388(94)00128-6.

[6]	J. I. Calvo, A. Bottino, G. Capannelli, and A. Hernández, "Comparison of liquid–liquid displacement porosimetry and scanning electron microscopy image analysis to characterise ultrafiltration track-etched membranes," Journal of Membrane Science, vol. 239, no. 2, pp. 189-197, 2004/08/15/ 2004, doi: https://doi.org/10.1016/j.memsci.2004.02.038.

[7]	N. A. Ochoa, P. Prádanos, L. Palacio, C. Pagliero, J. Marchese, and A. Hernández, "Pore size distributions based on AFM imaging and retention of multidisperse polymer solutes: Characterisation of polyethersulfone UF membranes with dopes containing different PVP," Journal of Membrane Science, vol. 187, no. 1, pp. 227-237, 2001/06/15/ 2001, doi: https://doi.org/10.1016/S0376-7388(01)00348-9.

[8]	H. Reingruber, A. Zankel, C. Mayrhofer, and P. Poelt, "A new in situ method for the characterization of membranes in a wet state in the environmental scanning electron microscope," Journal of Membrane Science, Article vol. 399-400, pp. 86-94, 2012, doi: 10.1016/j.memsci.2012.01.031.

[9]	M. B. Tanis-Kanbur, R. I. Peinador, X. Hu, J. I. Calvo, and J. W. Chew, "Membrane characterization via evapoporometry (EP) and liquid-liquid displacement porosimetry (LLDP) techniques," Journal of Membrane Science, vol. 586, pp. 248-258, 2019/09/15/ 2019, doi: https://doi.org/10.1016/j.memsci.2019.05.077.

[10]	M. Baklanov, K. P. Mogilnikov, V. Polovinkin, and F. Dultsev, Determination of pore size distribution in thin films by ellipsometric porosimetry. 2000, pp. 1385-1391.

[11]	H. D. Bale and P. W. Schmidt, "Small-Angle X-Ray-Scattering Investigation of Submicroscopic Porosity with Fractal Properties," Physical Review Letters, Article vol. 53, no. 6, pp. 596-599, 1984, doi: 10.1103/PhysRevLett.53.596.

[12]	T. C. Chilcott, M. Chan, L. Gaedt, T. Nantawisarakul, A. G. Fane, and H. G. L. Coster, "Electrical impedance spectroscopy characterisation of conducting membranes: I. Theory," Journal of Membrane Science, vol. 195, no. 2, pp. 153-167, 2002/01/31/ 2002, doi: https://doi.org/10.1016/S0376-7388(01)00541-5.

[13]	K. Darowicki and M. Szociński, "Local impedance spectroscopy of membranes," Journal of Membrane Science, Article vol. 303, no. 1-2, pp. 1-3, 2007, doi: 10.1016/j.memsci.2007.06.063.

[14]	J.-D. Jeon, S. J. Kim, and S.-Y. Kwak, "1H nuclear magnetic resonance (NMR) cryoporometry as a tool to determine the pore size distribution of ultrafiltration membranes," Journal of Membrane Science, vol. 309, no. 1, pp. 233-238, 2008/02/15/ 2008, doi: https://doi.org/10.1016/j.memsci.2007.10.034.

[15]	K. C. Khulbe, B. Kruczek, G. Chowdhury, S. Gagné, T. Matsuura, and S. P. Verma, "Characterization of membranes prepared from PPO by Raman scattering and atomic force microscopy," Journal of Membrane Science, vol. 111, no. 1, pp. 57-70, 1996/03/06/ 1996, doi: https://doi.org/10.1016/0376-7388(95)00276-6.

[16]	K. C. Khulbe and T. Matsuura, "Characterization of synthetic membranes by Raman spectroscopy, electron spin resonance, and atomic force microscopy; a review," Polymer, vol. 41, no. 5, pp. 1917-1935, 2000/03/01/ 2000, doi: https://doi.org/10.1016/S0032-3861(99)00359-6.

[17]	K. C. Khulbe, T. Matsuura, and C. Y. Feng, "Study on cellulose acetate membranes for reverse osmosis and polyethersulfone membranes for ultrafiltration by electron spin resonance technique," Desalination, vol. 148, no. 1, pp. 329-332, 2002/09/10/ 2002, doi: https://doi.org/10.1016/S0011-9164(02)00725-7.

[18]	K. C. Khulbe, T. Matsuura, G. Lamarche, A. M. Lamarche, C. Choi, and S. H. Noh, "Study of the structure of asymmetric cellulose acetate membranes for reverse osmosis using electron spin resonance (ESR) method," Polymer, vol. 42, no. 15, pp. 6479-6484, 2001/07/01/ 2001, doi: https://doi.org/10.1016/S0032-3861(01)00156-2.

[19]	R. Lamsal, S. G. Harroun, C. L. Brosseau, and G. A. Gagnon, "Use of surface enhanced Raman spectroscopy for studying fouling on nanofiltration membrane," Separation and Purification Technology, vol. 96, pp. 7-11, 2012/08/21/ 2012, doi: https://doi.org/10.1016/j.seppur.2012.05.019.

[20]	D. Nanda et al., "Characterization of fouled nanofiltration membranes using positron annihilation spectroscopy," Journal of Membrane Science, vol. 382, no. 1, pp. 124-134, 2011/10/15/ 2011, doi: https://doi.org/10.1016/j.memsci.2011.08.026.

[21]	P. S. Singh, P. Ray, Z. Xie, and M. Hoang, "Synchrotron SAXS to probe cross-linked network of polyamide ‘reverse osmosis’ and ‘nanofiltration’ membranes," Journal of Membrane Science, vol. 421-422, pp. 51-59, 2012/12/01/ 2012, doi: https://doi.org/10.1016/j.memsci.2012.06.029.

[22]	H.-K. Song, Y.-H. Jung, K.-H. Lee, and L. H. Dao, "Electrochemical impedance spectroscopy of porous electrodes: the effect of pore size distribution," Electrochimica Acta, vol. 44, no. 20, pp. 3513-3519, 1999/06/01/ 1999, doi: https://doi.org/10.1016/S0013-4686(99)00121-8.

[23]	K.-L. Tung et al., "Characterization of multilayer nanofiltration membranes using positron annihilation spectroscopy," Journal of Membrane Science, vol. 343, no. 1, pp. 147-156, 2009/11/01/ 2009, doi: https://doi.org/10.1016/j.memsci.2009.07.020.

[24]	M. B. Tanis-Kanbur, F. Zamani, W. B. Krantz, X. Hu, and J. W. Chew, "Adaptation of evapoporometry (EP) to characterize the continuous pores and interpore connectivity in polymeric membranes," Journal of Membrane Science, vol. 575, pp. 17-27, 2019/04/01/ 2019, doi: https://doi.org/10.1016/j.memsci.2018.12.050.

[25]	K. J. Kim et al., "A comparative study of techniques used for porous membrane characterization: pore characterization," Journal of Membrane Science, vol. 87, no. 1, pp. 35-46, 1994/02/23/ 1994, doi: https://doi.org/10.1016/0376-7388(93)E0044-E.

[26]	C. Zhao, X. Zhou, and Y. Yue, "Determination of pore size and pore size distribution on the surface of hollow-fiber filtration membranes: a review of methods," Desalination, vol. 129, no. 2, pp. 107-123, 2000/07/10/ 2000, doi: https://doi.org/10.1016/S0011-9164(00)00054-0.

[27]	M. A. Rahman, M. A. Mutalib, K. Li, and M. H. D. Othman, "Chapter 10 - Pore Size Measurements and Distribution for Ceramic Membranes," in Membrane Characterization, N. Hilal, A. F. Ismail, T. Matsuura, and D. Oatley-Radcliffe Eds.: Elsevier, 2017, pp. 183-198.

[28]	K. C. Khulbe and T. Matsuura, "Synthetic membrane characterisation – a review: part I," Membrane Technology, Article vol. 2017, no. 7, pp. 7-12, 2017, doi: 10.1016/S0958-2118(17)30135-0.

[29]	K. C. Khulbe, T. Matsuura, and C. Y. Feng, "Synthetic membrane characterisation – a review: part II," Membrane Technology, Article vol. 2017, no. 8, pp. 7-12, 2017, doi: 10.1016/S0958-2118(17)30171-4.

[30]	K. L. Tung, K. S. Chang, T. T. Wu, N. J. Lin, K. R. Lee, and J. Y. Lai, "Recent advances in the characterization of membrane morphology," Current Opinion in Chemical Engineering, Review vol. 4, pp. 121-127, 2014, doi: 10.1016/j.coche.2014.03.002.

[31]	N. Hilal, A. F. Ismail, T. Matsuura, and D. Oatley-Radcliffe, "Membrane Characterization," in Membrane Characterization, N. Hilal, A. F. Ismail, T. Matsuura, and D. Oatley-Radcliffe Eds.: Elsevier, 2017, pp. xix-xxi.

[32]	K. Friess, L. Bartovská, K. Pilnáček, O. Vopička, A. Randová, and M. Lanč, "Physical Chemistry Characterization of Polymeric Membranes," 2017.

[33]	D. J. Johnson, D. L. Oatley-Radcliffe, and N. Hilal, "State of the art review on membrane surface characterisation: Visualisation, verification and quantification of membrane properties," Desalination, Review vol. 434, pp. 12-36, 2018, doi: 10.1016/j.desal.2017.03.023.

[34]	W. B. Krantz, A. R. Greenberg, E. Kujundzic, A. Yeo, and S. S. Hosseini, "Evapoporometry: A novel technique for determining the pore-size distribution of membranes," Journal of Membrane Science, vol. 438, pp. 153-166, Jul 2013, doi: 10.1016/j.memsci.2013.03.045.

[35]	M. Brun, A. Lallemand, J.-F. Quinson, and C. Eyraud, "A new method for the simultaneous determination of the size and shape of pores: the thermoporometry," Thermochimica Acta, vol. 21, no. 1, pp. 59-88, 1977/10/01/ 1977, doi: https://doi.org/10.1016/0040-6031(77)85122-8.

[36]	T. Fujioka, N. Oshima, R. Suzuki, W. E. Price, and L. D. Nghiem, "Probing the internal structure of reverse osmosis membranes by positron annihilation spectroscopy: Gaining more insight into the transport of water and small solutes," Journal of Membrane Science, Review vol. 486, pp. 106-118, 2015, doi: 10.1016/j.memsci.2015.02.007.

[37]	R. W. Siegel, "Positron Annihilation Spectroscopy," Annual Review of Materials Science, vol. 10, no. 1, pp. 393-425, 1980, doi: 10.1146/annurev.ms.10.080180.002141.

[38]	Y. C. Jean, P. E. Mallon, and D. M. Schrader, "Introduction to Positron and Positronium Chemistry," in Principles and Applications of Positron and Positronium Chemistry, pp. 1-15.

[39]	W. J. Lau, A. F. Ismail, P. S. Goh, N. Hilal, and B. S. Ooi, "Characterization methods of thin film composite nanofiltration membranes," Separation and Purification Reviews, Review vol. 44, no. 2, pp. 135-156, 2015, doi: 10.1080/15422119.2014.882355.

[40]	M. Qasim, M. Badrelzaman, N. N. Darwish, N. A. Darwish, and N. Hilal, "Reverse osmosis desalination: A state-of-the-art review," Desalination, Review vol. 459, pp. 59-104, 2019, doi: 10.1016/j.desal.2019.02.008.

[41]	H. H. Yin, Y. Zejie, M. Weitao, and Z. Daming, "A Review of Studies of Polymeric Membranes by Positron Annihilation Lifetime Spectroscopy," Plasma Science and Technology, vol. 7, no. 5, pp. 3062-3064, 2005/10 2005, doi: 10.1088/1009-0630/7/5/020.

[42]	M. Eldrup, D. Lightbody, and J. N. Sherwood, "The temperature dependence of positron lifetimes in solid pivalic acid," Chemical Physics, vol. 63, no. 1, pp. 51-58, 1981/12/01/ 1981, doi: https://doi.org/10.1016/0301-0104(81)80307-2.

[43]	S. J. Tao, "Positronium annihilation in molecular substances," The Journal of Chemical Physics, Article vol. 56, no. 11, pp. 5499-5510, 1972, doi: 10.1063/1.1677067.

[44]	D. W. Gidley, H.-G. Peng, and R. S. Vallery, "POSITRON ANNIHILATION AS A METHOD TO CHARACTERIZE POROUS MATERIALS," Annual Review of Materials Research, vol. 36, no. 1, pp. 49-79, 2006, doi: 10.1146/annurev.matsci.36.111904.135144.

[45]	P. Dumas and V. Humblot, "Chapter 6 - Synchrotron infrared interface science," in Biointerface Characterization by Advanced IR Spectroscopy, C. M. Pradier and Y. J. Chabal Eds. Amsterdam: Elsevier, 2011, pp. 145-166.

[46]	S. L. Hulbert and G. P. Williams, "1 - SYNCHROTRON RADIATION SOURCES," in Vacuum Ultraviolet Spectroscopy, J. A. R. Samson and D. L. Ederer Eds. Burlington: Academic Press, 2000, pp. 1-25.

[47]	B. A. Cunningham, W. Bras, L. J. Lis, and P. J. Quinn, "Synchrotron X-ray studies of lipids and membranes: a critique," Journal of Biochemical and Biophysical Methods, vol. 29, no. 2, pp. 87-111, 1994/09/01/ 1994, doi: https://doi.org/10.1016/0165-022X(94)90046-9.

[48]	F. C. Adams, "X-Ray Absorption and Diffraction | Overview," in Encyclopedia of Analytical Science (Third Edition), P. Worsfold, C. Poole, A. Townshend, and M. Miró Eds. Oxford: Academic Press, 2019, pp. 391-403.

[49]	H. Bechhold, "The permeability of ultrafilters," Z. Phys. Chem., vol. 64, pp. 328-342, 1908.

[50]	J. I. Calvo, "Membrane characterization by porosimetric techniques," presented at the XXIV Summer School on Membranes, Genoa, Italy, 2007.

[51]	R. G., "Method for Determining the Bubble Point or the Largest Pore of Membranes or Filter Materials," US, 1988.

[52]	 G. Reichelt, "Bubble point measurements on large membrane areas," in Polymeric Materials Science and Engineering, Proceedings of the ACS Division of Polymeric Materials Science and Engineering, 1989, vol. 61, pp. 787-788.

[53]	G. Reichelt, "Bubble point measurements on large areas of microporous membranes," Journal of membrane science, vol. 60, no. 2-3, pp. 253-259, 1991.

[54]	S. K. Bhatia and J. L. Smith, "Comparative study of bubble point method and mercury intrusion porosimetry techniques for characterizing the pore-size distribution of geotextiles," Geotextiles and Geomembranes, vol. 13, no. 10, pp. 679-702, 1994.

[55]	S. Bhatia and J. Smith, "Application of the bubble point method to the characterization of the pore-size distribution of geotextiles," Geotechnical Testing Journal, vol. 18, no. 1, pp. 94-105, 1995.

[56]	A. Hernández, J. Calvo, P. Prádanos, and F. Tejerina, "Pore size distributions in microporous membranes. A critical analysis of the bubble point extended method," Journal of Membrane Science, vol. 112, no. 1, pp. 1-12, 1996.

[57]	E. Jakobs and W. Koros, "Ceramic membrane characterization via the bubble point technique," Journal of Membrane Science, vol. 124, no. 2, pp. 149-159, 1997.

[58]	J. Yu, X. Hu, and Y. Huang, "A modification of the bubble-point method to determine the pore-mouth size distribution of porous materials," Separation and purification technology, vol. 70, no. 3, pp. 314-319, 2010.

[59]	V. Nassehi, D. B. Das, I. M. Shigidi, and R. J. Wakeman, "Numerical analyses of bubble point tests used for membrane characterisation: model development and experimental validation," Asia‐Pacific Journal of Chemical Engineering, vol. 6, no. 6, pp. 850-862, 2011.

[60]	D. Hopkinson, M. Zeh, and D. Luebke, "The bubble point of supported ionic liquid membranes using flat sheet supports," Journal of membrane science, vol. 468, pp. 155-162, 2014.

[61]	F. Erbe, "Die Bestimmung der Porenverteilung nach ihrer Größe in Filtern und Ultrafiltern," Kolloid-Zeitschrift, Article vol. 63, no. 3, pp. 277-285, 1933, doi: 10.1007/BF01422935.

[62]	S. P. Nunes and K. V. Peinemann, "Ultrafiltration membranes from PVDF/PMMA blends," Journal of Membrane Science, vol. 73, no. 1, pp. 25-35, 1992/10/02/ 1992, doi: https://doi.org/10.1016/0376-7388(92)80183-K.

[63]	 H. Steinhauser, Scholz, H., Hübner, A., Ellinghorst, G. , "The role of supports for pervaporation composite membranes," in Proc. ICOM 90, Chicago, 1990, vol. 1.

[64]	P. Schneider and P. Uchytil, "Liquid expulsion permporometry for characterization of porous membranes," Journal of Membrane Science, vol. 95, no. 1, pp. 29-38, 1994/10/10/ 1994, doi: https://doi.org/10.1016/0376-7388(94)85026-7.

[65]	A. Hernández, J. I. Calvo, P. Prádanos, and F. Tejerina, "Pore size distributions in microporous membranes. A critical analysis of the bubble point extended method," Journal of Membrane Science, Article vol. 112, no. 1, pp. 1-12, 1996, doi: 10.1016/0376-7388(95)00025-9.

[66]	P. Shao, R. Y. M. Huang, X. Feng, and W. Anderson, "Gas-Liquid Displacement Method for Estimating Membrane Pore-Size Distributions," AIChE Journal, Article vol. 50, no. 3, pp. 557-565, 2004, doi: 10.1002/aic.10050.

[67]	Standard Test Method for Pore Size Characteristics of Membrane Filters Using Automated Liquid Porosimeter, A. E1294-89(1999), West Conshohocken, PA 1999.

[68]	Standard Test Methods for Pore Size Characteristics of Membrane Filters by Bubble Point and Mean Flow Pore Test, A. F316-03(2011), West Conshohocken, PA, 2011.

[69]	A. Shrestha, "Characterization of porous membranes via porometry," Master Thesis, Mechanical Engineering, University of Colorado at Boulder, 2012.

[70]	J. I. Calvo, A. Hernández, P. Prádanos, L. Martínez, and W. R. Bowen, "Pore size distributions in microporous membranes II. Bulk characterization of track-Etched filters by air porometry and mercury porosimetry," Journal of Colloid And Interface Science, Article vol. 176, no. 2, pp. 467-478, 1995, doi: 10.1006/jcis.1995.9944.

[71]	R. Gopal, S. Kaur, Z. Ma, C. Chan, S. Ramakrishna, and T. Matsuura, "Electrospun nanofibrous filtration membrane," Journal of Membrane Science, vol. 281, no. 1, pp. 581-586, 2006/09/15/ 2006, doi: https://doi.org/10.1016/j.memsci.2006.04.026.

[72]	D. Li, M. W. Frey, and Y. L. Joo, "Characterization of nanofibrous membranes with capillary flow porometry," Journal of Membrane Science, vol. 286, no. 1, pp. 104-114, 2006/12/15/ 2006, doi: https://doi.org/10.1016/j.memsci.2006.09.020.

[73]	W. Piątkiewicz, S. Rosiński, D. Lewińska, J. Bukowski, and W. Judycki, "Determination of pore size distribution in hollow fibre membranes," Journal of Membrane Science, vol. 153, no. 1, pp. 91-102, 1999/02/03/ 1999, doi: https://doi.org/10.1016/S0376-7388(98)00243-9.

[74]	X. Liu, Z. Qiankun, H. Zhang, Y. Jiang, and Y. He, "Development and characterization of microporous Ti3SiC2 ceramic membranes for filtration of microorganisms," Journal of Materials Science, vol. 51, 11/16 2015, doi: 10.1007/s10853-015-9572-1.

[75]	C. Wang, L. Ling, Y. Huang, Y. Yao, and Q. Song, "Decoration of porous ceramic substrate with pencil for enhanced gas separation performance of carbon membrane," Carbon, vol. 84, pp. 151-159, 2015/04/01/ 2015, doi: https://doi.org/10.1016/j.carbon.2014.12.003.

[76]	D. Y. Wenten IG, "Metode dan alat bubble point untuk karakterisasi membran mikrofiltrasi " Indonesia, 2002.

[77]	R. Mourhatch, T. T. Tsotsis, and M. Sahimi, "Determination of the true pore size distribution by flow permporometry experiments: An invasion percolation model," Journal of Membrane Science, vol. 367, no. 1, pp. 55-62, 2011/02/01/ 2011, doi: https://doi.org/10.1016/j.memsci.2010.10.042.

[78]	C. Herrero, P. Prádanos, J. I. Calvo, F. Tejerina, and A. Hernández, "Flux Decline in Protein Microfiltration: Influence of Operative Parameters," Journal of Colloid and Interface Science, vol. 187, no. 2, pp. 344-351, 1997/03/15/ 1997, doi: https://doi.org/10.1006/jcis.1996.4662.

[79]	J. Jacob, P. Prádanos, J. I. Calvo, A. Hernández, and G. Jonsson, "Fouling kinetics and associated dynamics of structural modifications," Colloids and Surfaces A: Physicochemical and Engineering Aspects, vol. 138, no. 2, pp. 173-183, 1998/07/15/ 1998, doi: https://doi.org/10.1016/S0927-7757(97)00082-4.

[80]	M. Mietton-Peuchot, C. Condat, and T. Courtois, "Use of gas-liquid porometry measurements for selection of microfiltration membranes," Journal of Membrane Science, vol. 133, no. 1, pp. 73-82, 1997/09/17/ 1997, doi: https://doi.org/10.1016/S0376-7388(97)00091-4.

[81]	C. M. Gribble et al., "Porometry, porosimetry, image analysis and void network modelling in the study of the pore-level properties of filters," Chemical Engineering Science, vol. 66, no. 16, pp. 3701-3709, 2011/08/15/ 2011, doi: https://doi.org/10.1016/j.ces.2011.05.013.

[82]	M. A. Islam, M. S. Hossain, and M. Ulbricht, "Model-dependent analysis of gas flow/pore dewetting data for microfiltration membranes," Journal of Membrane Science, vol. 533, pp. 351-363, 2017/07/01/ 2017, doi: https://doi.org/10.1016/j.memsci.2017.03.012.

[83]	M. A. Islam and M. Ulbricht, "Microfiltration membrane characterization by gas-liquid displacement porometry: Matching experimental pore number distribution with liquid permeability and bulk porosity," Journal of Membrane Science, vol. 569, pp. 104-116, 2019/01/01/ 2019, doi: https://doi.org/10.1016/j.memsci.2018.09.030.

[84]	S. Ramaswamy, A. Greenberg, and M. Peterson, "Non-invasive measurement of membrane morphology via UFDR: Pore-size characterization," Journal of Membrane Science, vol. 239, pp. 143-154, 08/01 2004, doi: 10.1016/j.memsci.2003.08.030.

[85]	S. S Manickam and J. R. McCutcheon, "Characterization of polymeric nonwovens using porosimetry, porometry and X-ray computed tomography," Journal of Membrane Science, vol. 407-408, pp. 108-115, 2012/07/15/ 2012, doi: https://doi.org/10.1016/j.memsci.2012.03.022.

[86]	B. Miller and I. Tyomkin, "Liquid Porosimetry: New Methodology and Applications," Journal of Colloid and Interface Science, vol. 162, no. 1, pp. 163-170, 1994/01/01/ 1994, doi: https://doi.org/10.1006/jcis.1994.1021.

[87]	A. Jen and K. Gupta, "Characterization of pore structure of filter media," Fluid - Particle Separation Journal, vol. 14, 12/01 2002.

[88]	J.-G. Lee et al., "Theoretical modeling and experimental validation of transport and separation properties of carbon nanotube electrospun membrane distillation," Journal of Membrane Science, vol. 526, pp. 395-408, 2017/03/15/ 2017, doi: https://doi.org/10.1016/j.memsci.2016.12.045.

[89]	 A. H. José Ignacio Calvo, Laura Palacio, Pedro Prádanos, Matteo Ailuno, Aldo, Bottino, G.  Carniglia, Antonio Comité, A. Jezowska, M. Pagliero and Rafaella Firpo, "Characterization of porous PTFE membranes and their potential application in osmotic distillation," in Euromembrane Valencia, Spain, 2018.

[90]	J. T. Gostick, M. A. Ioannidis, M. W. Fowler, and M. D. Pritzker, "Direct measurement of the capillary pressure characteristics of water–air–gas diffusion layer systems for PEM fuel cells," Electrochemistry Communications, vol. 10, no. 10, pp. 1520-1523, 2008/10/01/ 2008, doi: https://doi.org/10.1016/j.elecom.2008.08.008.

[91]	H. E. Kolb, R. Schmitt, A. Dittler, and G. Kasper, "On the accuracy of capillary flow porometry for fibrous filter media," Separation and Purification Technology, vol. 199, pp. 198-205, 2018/06/30/ 2018, doi: https://doi.org/10.1016/j.seppur.2018.01.024.

[92]	D. Dollimore and G. R. Heal, "An improved method for the calculation of pore size distribution from adsorption data," Journal of Applied Chemistry, vol. 14, no. 3, pp. 109-114, 1964/03/01 1964, doi: 10.1002/jctb.5010140302.

[93]	G. Capannelli, F. Vigo, and S. Munari, "Ultrafiltration membranes — characterization methods," Journal of Membrane Science, vol. 15, no. 3, pp. 289-313, 1983/10/01/ 1983, doi: https://doi.org/10.1016/S0376-7388(00)82305-4.

[94]	R. E. Kesting, Synthetic polymeric membranes: a structural perspective. Wiley, 1985.

[95]	S. Munari, A. Bottino, G. Capannelli, and P. Moretti, "Membrane morphology and transport properties," Desalination, vol. 53, no. 1, pp. 11-23, 1985/01/01/ 1985, doi: https://doi.org/10.1016/0011-9164(85)85049-9.

[96]	S. N. P. Grabar, "Sur le diamètre des pores des membranes en collodion utilisées en ultrafiltration," J. Chim. Phys, vol. 33 no. (1936) pp. 721–741.

[97]	R. I. Peinador, J. I. Calvo, P. Prádanos, L. Palacio, and A. Hernández, "Characterisation of polymeric UF membranes by liquid–liquid displacement porosimetry," Journal of Membrane Science, vol. 348, no. 1, pp. 238-244, 2010/02/15/ 2010, doi: https://doi.org/10.1016/j.memsci.2009.11.008.

[98]	K. S. McGuire, K. W. Lawson, and D. R. Lloyd, "Pore size distribution determination from liquid permeation through microporous membranes," Journal of Membrane Science, vol. 99, no. 2, pp. 127-137, 1995/02/28/ 1995, doi: https://doi.org/10.1016/0376-7388(94)00209-H.

[99]	A. J. Gijsbertsen-Abrahamse, R. M. Boom, and A. van der Padt, "Why liquid displacement methods are sometimes wrong in estimating the pore-size distribution," AIChE Journal, vol. 50, no. 7, pp. 1364-1371, 2004, doi: 10.1002/aic.10124.

[100]	K. R. Morison, "A comparison of liquid–liquid porosimetry equations for evaluation of pore size distribution," Journal of Membrane Science, vol. 325, no. 1, pp. 301-310, 2008/11/15/ 2008, doi: https://doi.org/10.1016/j.memsci.2008.07.042.

[101]	M. C. Almécija et al., "Analysis of cleaning protocols in ceramic membranes by liquid–liquid displacement porosimetry," Desalination, vol. 245, no. 1, pp. 541-545, 2009/09/15/ 2009, doi: https://doi.org/10.1016/j.desal.2009.02.018.

[102]	E. Anton, J. I. Calvo, J. R. Alvarez, A. Hernandez, and S. Luque, "Fitting approach to liquid-liquid displacement porosimetry based on the log-normal pore size distribution," Journal of Membrane Science, vol. 470, pp. 219-228, Nov 2014, doi: 10.1016/j.memsci.2014.07.035.

[103]	J. I. Calvo, A. Bottino, G. Capannelli, and A. Hernández, "Pore size distribution of ceramic UF membranes by liquid–liquid displacement porosimetry," Journal of Membrane Science, vol. 310, no. 1, pp. 531-538, 2008/03/05/ 2008, doi: https://doi.org/10.1016/j.memsci.2007.11.035.

[104]	J. I. Calvo et al., "Porosimetric characterization of polysulfone ultrafiltration membranes by image analysis and liquid–liquid displacement technique," Desalination, vol. 357, pp. 84-92, 2015/02/02/ 2015, doi: https://doi.org/10.1016/j.desal.2014.11.012.

[105]	J. I. Calvo et al., "Liquid–liquid displacement porometry to estimate the molecular weight cut-off of ultrafiltration membranes," Desalination, vol. 268, no. 1, pp. 174-181, 2011/03/01/ 2011, doi: https://doi.org/10.1016/j.desal.2010.10.016.

[106]	P. Carretero et al., "Liquid–liquid displacement porosimetry applied to several MF and UF membranes," Desalination, vol. 327, no. Supplement C, pp. 14-23, 2013/10/15/ 2013, doi: https://doi.org/10.1016/j.desal.2013.08.001.

[107]	R. I. Peinador, J. I. Calvo, K. ToVinh, V. Thom, P. Prádanos, and A. Hernández, "Liquid–liquid displacement porosimetry for the characterization of virus retentive membranes," Journal of Membrane Science, vol. 372, no. 1, pp. 366-372, 2011/04/15/ 2011, doi: https://doi.org/10.1016/j.memsci.2011.02.022.

[108]	A. Bottino, G. Capannelli, A. Grosso, O. Monticelli, and M. Nicchia, "Porosimetric Characterization of Inorganic Membranes," Separation Science and Technology, vol. 29, no. 8, pp. 985-999, 1994/05/01 1994, doi: 10.1080/01496399408005612.

[109]	A. Bottino, G. Capannelli, P. Petit-bon, N. Cao, M. Pegoraro, and G. Zoia, "Pore Size and Pore-Size Distribution in Microfiltration Membranes," Separation Science and Technology, vol. 26, no. 10-11, pp. 1315-1327, 1991/10/01 1991, doi: 10.1080/01496399108050534.

[110]	G. Capannelli, I. Becchi, A. Bottino, P. Moretti, and S. Munari, "Computer Driven Porosimeter for Ultrafiltration Membranes," in Studies in Surface Science and Catalysis, vol. 39, K. K. Unger, J. Rouquerol, K. S. W. Sing, and H. Kral Eds.: Elsevier, 1988, pp. 283-294.

[111]	T. Courtois, "Dispositif de mesure de caractéristiques liées a la porosité des media poreux," France, 2003.

[112]	S. Munari, A. Bottino, P. Moretti, G. Capannelli, and I. Becchi, "Permoporometric study on ultrafiltration membranes," Journal of Membrane Science, vol. 41, pp. 69-86, 1989/02/15/ 1989, doi: https://doi.org/10.1016/S0376-7388(00)82392-3.

[113]	E. S. Tarleton, J. P. Robinson, and M. Salman, "Solvent-induced swelling of membranes — Measurements and influence in nanofiltration," Journal of Membrane Science, vol. 280, no. 1, pp. 442-451, 2006/09/01/ 2006, doi: https://doi.org/10.1016/j.memsci.2006.01.050.

[114]	L. Germic, K. Ebert, R. H. B. Bouma, Z. Borneman, M. H. V. Mulder, and H. Strathmann, "Characterization of polyacrylonitrile ultrafiltration membranes," Journal of Membrane Science, vol. 132, no. 1, pp. 131-145, 1997/08/20/ 1997, doi: https://doi.org/10.1016/S0376-7388(97)00058-6.

[115]	M. W. Phillips and A. J. DiLeo, "A Validatible Porosimetric Technique for Verifying the Integrity of Virus-Retentive Membranes," Biologicals, vol. 24, no. 3, pp. 243-253, 1996/09/01/ 1996, doi: https://doi.org/10.1006/biol.1996.0033.

[116]	S. M. Snyder, K. D. Cole, and D. C. Szlag, "Phase compositions, viscosities, and densities for aqueous two-phase systems composed of polyethylene glycol and various salts at 25 .degree.C," Journal of Chemical & Engineering Data, vol. 37, no. 2, pp. 268-274, 1992/04/01 1992, doi: 10.1021/je00006a036.

[117]	P. A. Albertsson, "Partition of cell particles and macromolecules " Advances in protein chemistry, vol. 24, p. 309, 1986.

[118]	E. Atefi, J. A. Mann, and H. Tavana, "Ultralow Interfacial Tensions of Aqueous Two-Phase Systems Measured Using Drop Shape," Langmuir, vol. 30, no. 32, pp. 9691-9699, 2014/08/19 2014, doi: 10.1021/la500930x.

[119]	D. Forciniti, C. K. Hall, and M. R. Kula, "Influence of polymer molecular weight and temperature on phase composition in aqueous two-phase systems," Fluid Phase Equilibria, vol. 61, no. 3, pp. 243-262, 1991/10/01/ 1991, doi: https://doi.org/10.1016/0378-3812(91)80002-D.

[120]	D. Forciniti, C. K. Hall, and M. R. Kula, "Analysis of polymer molecular weight distributions in aqueous two-phase systems," Journal of Biotechnology, vol. 20, no. 2, pp. 151-161, 1991/09/01/ 1991, doi: https://doi.org/10.1016/0168-1656(91)90224-J.

[121]	B. Mattiasson, "Applications of aqueous two-phase systems in biotechnology," Trends in Biotechnology, vol. 1, no. 1, pp. 16-20, 1983/03/01/ 1983, doi: https://doi.org/10.1016/0167-7799(83)90021-5.

[122]	P. Cheng, D. Li, L. Boruvka, Y. Rotenberg, and A. W. Neumann, "Automation of axisymmetric drop shape analysis for measurements of interfacial tensions and contact angles," Colloids and Surfaces, vol. 43, no. 2, pp. 151-167, 1990/01/01/ 1990, doi: https://doi.org/10.1016/0166-6622(90)80286-D.

[123]	H. Cárdenas, M. Cartes, and A. Mejía, "Atmospheric densities and interfacial tensions for 1-alkanol (1-butanol to 1-octanol)+water and ether (MTBE, ETBE, DIPE, TAME and THP)+water demixed mixtures," Fluid Phase Equilibria, vol. 396, pp. 88-97, 2015/06/25/ 2015, doi: https://doi.org/10.1016/j.fluid.2015.03.040.

[124]	J. Viades-Trejo and J. Gracia-Fadrique, "Spinning drop method: From Young–Laplace to Vonnegut," Colloids and Surfaces A: Physicochemical and Engineering Aspects, vol. 302, no. 1, pp. 549-552, 2007/07/20/ 2007, doi: https://doi.org/10.1016/j.colsurfa.2007.03.033.

[125]	S. B. R. Peinador, "La caractérisation des membranes d'Ultrafiltration et Microfiltration par les techniques de pénétration FluideFluide (Gaz-Liquide et Liquide-Liquide)," in "Rapport de stage de fin d’études," Université toulouse III – Paul Sabatier, 2019.

[126]	S. Giglia, D. Bohonak, P. Greenhalgh, and A. Leahy, "Measurement of pore size distribution and prediction of membrane filter virus retention using liquid–liquid porometry," Journal of Membrane Science, vol. 476, pp. 399-409, 2015/02/15/ 2015, doi: https://doi.org/10.1016/j.memsci.2014.11.053.

[127]	A. A. Liabastre and C. Orr, "An evaluation of pore structure by mercury penetration," Journal of Colloid and Interface Science, vol. 64, no. 1, pp. 1-18, 1978/03/15/ 1978, doi: https://doi.org/10.1016/0021-9797(78)90329-6.

[128]	S. Lowell, Powder Surface Area and Porosity, Third Edition. ed. (Powder Technology Series, 2). Dordrecht: Springer Netherlands, 1991.

[129]	E. W. Washburn, "Note on a Method of Determining the Distribution of Pore Sizes in a Porous Material," (in eng), Proc Natl Acad Sci U S A, vol. 7, no. 4, pp. 115-116, 1921/04// 1921, doi: 10.1073/pnas.7.4.115.

[130]	E. Honold and E. L. Skau, "Application of Mercury-Intrusion Method for Determination of Pore-Size Distribution to Membrane Filters," Science, vol. 120, no. 3124, pp. 805-806, 1954, doi: 10.1126/science.120.3124.805.

[131]	F. Théron, E. Lys, A. Joubert, F. Bertrand, and L. Le Coq, "Characterization of the porous structure of a non-woven fibrous medium for air filtration at local and global scales using porosimetry and X-ray micro-tomography," Powder Technology, vol. 320, pp. 295-303, 2017/10/01/ 2017, doi: https://doi.org/10.1016/j.powtec.2017.07.020.

[132]	P. Webb, "An Introduction To The Physical Characterization of Materials by Mercury Intrusion Porosimetry with Emphasis On Reduction And Presentation of Experimental Data 2 CONTENTS," Pharm Online, 01/01 2001.

[133]	C. Volzone and N. Zagorodny, "Mercury intrusion porosimetry (MIP) study of archaeological pottery from Hualfin Valley, Catamarca, Argentina," Applied Clay Science, vol. 91-92, pp. 12-15, 2014/04/01/ 2014, doi: https://doi.org/10.1016/j.clay.2014.02.002.

[134]	I. Zake-Tiluga, V. Svinka, R. Svinka, B. Zierath, P. Greil, and T. Fey, "Thermal conductivity and microstructure characterisation of lightweight alumina and alumina–mullite ceramics," Journal of the European Ceramic Society, vol. 36, no. 6, pp. 1469-1477, 2016, doi: 10.1016/j.jeurceramsoc.2015.12.026.

[135]	L. Palacioa et al., "Ceramic membranes from Moroccan natural clay and phosphate for industrial water treatment," Desalination, vol. 245, no. 1-3, pp. 501-507, Sep 2009, doi: 10.1016/j.desal.2009.02.014.

[136]	B. F. K. Kingsbury and K. Li, "A morphological study of ceramic hollow fibre membranes," Journal of Membrane Science, vol. 328, no. 1, pp. 134-140, 2009/02/20/ 2009, doi: https://doi.org/10.1016/j.memsci.2008.11.050.

[137]	S. Koonaphapdeelert, Z. Wu, and K. Li, "Carbon dioxide stripping in ceramic hollow fibre membrane contactors," Chemical Engineering Science, vol. 64, no. 1, pp. 1-8, 2009/01/01/ 2009, doi: https://doi.org/10.1016/j.ces.2008.09.010.

[138]	Q. Liu, T. Xue, L. Yang, X. Hu, and H. Du, "Controllable synthesis of hierarchical porous mullite fiber network for gas filtration," Journal of the European Ceramic Society, vol. 36, no. 7, pp. 1691-1697, 2016/06/01/ 2016, doi: https://doi.org/10.1016/j.jeurceramsoc.2016.01.037.

[139]	M. H. D. Othman, N. Droushiotis, Z. Wu, G. Kelsall, and K. Li, "Dual-layer hollow fibres with different anode structures for micro-tubular solid oxide fuel cells," Journal of Power Sources, vol. 205, pp. 272-280, 2012/05/01/ 2012, doi: https://doi.org/10.1016/j.jpowsour.2012.01.002.

[140]	M. A. Rahman, F. R. García-García, M. D. I. Hatim, B. F. K. Kingsbury, and K. Li, "Development of a catalytic hollow fibre membrane micro-reactor for high purity H2 production," Journal of Membrane Science, vol. 368, no. 1, pp. 116-123, 2011/02/15/ 2011, doi: https://doi.org/10.1016/j.memsci.2010.11.025.

[141]	H. L. Ritter and L. C. Drake, "Pressure Porosimeter and Determination of Complete Macropore-Size Distributions. Pressure Porosimeter and Determination of Complete Macropore-Size Distributions," Industrial & Engineering Chemistry Analytical Edition, vol. 17, no. 12, pp. 782-786, 1945/12/01 1945, doi: 10.1021/i560148a013.

[142]	H. M. Rootare, "A Review of Mercury Porosimetry," in Advanced Experimental Techniques in Powder Metallurgy: Based on a Symposium on Advanced Experimental Techniques in Powder Metallurgy sponsored by the Institute of Metals Division, Powder Metallurgy Committee, held at the Spring Meeting of The Metallurgical Society of AIME in Pittsburgh, Pennsylvania, May 1969, J. S. Hirschhorn and K. H. Roll Eds. Boston, MA: Springer US, 1970, pp. 225-252.

[143]	C. A. León y León, "New perspectives in mercury porosimetry," Advances in Colloid and Interface Science, vol. 76-77, pp. 341-372, 1998/07/01/ 1998, doi: https://doi.org/10.1016/S0001-8686(98)00052-9.

[144]	A. Barrera et al., "Influence of the type of sepiolite on the modification of the pore-size distribution in γ-Al2O3 supports," Applied Clay Science, vol. 42, no. 3, pp. 415-421, 2009/01/01/ 2009, doi: https://doi.org/10.1016/j.clay.2008.04.009.

[145]	N. van Garderen et al., "Pore analyses of highly porous diatomite and clay based materials for fluidized bed reactors," Microporous and Mesoporous Materials, vol. 151, pp. 255–263, 03/01 2012, doi: 10.1016/j.micromeso.2011.10.028.

[146]	E. P. Favvas et al., "Characterization of carbonate rocks by combination of scattering, porosimetry and permeability techniques," Microporous and Mesoporous Materials, vol. 120, no. 1, pp. 109-114, 2009/04/01/ 2009, doi: https://doi.org/10.1016/j.micromeso.2008.09.015.

[147]	Y. Yao and D. Liu, "Comparison of low-field NMR and mercury intrusion porosimetry in characterizing pore size distributions of coals," Fuel, vol. 95, pp. 152-158, 2012/05/01/ 2012, doi: https://doi.org/10.1016/j.fuel.2011.12.039.

[148]	R. Zhang et al., "On the effect of mesoporosity of FAU Y zeolites in the liquid-phase catalysis," Microporous and Mesoporous Materials, vol. 278, pp. 297-306, 2019/04/01/ 2019, doi: https://doi.org/10.1016/j.micromeso.2018.12.003.

[149]	A. M. El-Kady and A. F. Ali, "Fabrication and characterization of ZnO modified bioactive glass nanoparticles," Ceramics International, vol. 38, no. 2, pp. 1195-1204, 2012/03/01/ 2012, doi: https://doi.org/10.1016/j.ceramint.2011.07.069.

[150]	J. M. Manso, Á. Rodriguez, Á. Aragón, and J. J. Gonzalez, "The durability of masonry mortars made with ladle furnace slag," Construction and Building Materials, vol. 25, no. 8, pp. 3508-3519, 2011/08/01/ 2011, doi: https://doi.org/10.1016/j.conbuildmat.2011.03.044.

[151]	S. P. Rigby and K. J. Edler, "The Influence of Mercury Contact Angle, Surface Tension, and Retraction Mechanism on the Interpretation of Mercury Porosimetry Data," Journal of Colloid and Interface Science, vol. 250, no. 1, pp. 175-190, 2002/06/01/ 2002, doi: https://doi.org/10.1006/jcis.2002.8286.

[152]	T. H. Sasakawa, Yasuhiro; Takami, Norio, "Electrode, secondary battery, battery pack, and vehicle," 2019.

[153]	V. Claude, J. G. Mahy, F. Micheli, J. Geens, and S. D. Lambert, "Sol-gel Ni/γ-Al2O3 material as secondary catalyst for toluene reforming: Tailoring the γ-Al2O3 substrate with stearic acid," Microporous and Mesoporous Materials, vol. 291, p. 109681, 2020/01/01/ 2020, doi: https://doi.org/10.1016/j.micromeso.2019.109681.

[154]	J. L. Amorós, A. Moreno, E. Blasco, J. J. Pérez, S. Navarro, and S. Reverter, "Inkjet technology for ceramic products. Influence of some process variables on ink penetration," International Journal of Applied Ceramic Technology, vol. 16, no. 6, pp. 2153-2160, 2019, doi: 10.1111/ijac.13342.

[155]	I. K. Park et al., "Three-dimensionally interconnected titanium foam anode for an energy-efficient zero gap-type chlor-alkali electrolyzer," International Journal of Hydrogen Energy, vol. 44, no. 31, pp. 16079-16086, 2019/06/21/ 2019, doi: https://doi.org/10.1016/j.ijhydene.2019.04.230.

[156]	S. Kerdi, A. Qamar, A. Alpatova, and N. Ghaffour, "An in-situ technique for the direct structural characterization of biofouling in membrane filtration," Journal of Membrane Science, vol. 583, pp. 81-92, 2019/08/01/ 2019, doi: https://doi.org/10.1016/j.memsci.2019.04.051.

[157]	A. Kosior, M. Antosova, R. Fáber, L. Villain, and M. Polakovič, "Single-component adsorption of proteins on a cellulose membrane with the phenyl ligand for hydrophobic interaction chromatography," Journal of Membrane Science, vol. 442, pp. 216–224, 09/01 2013, doi: 10.1016/j.memsci.2013.04.013.

[158]	H. J. Lee and J. H. Park, "Effect of hydrophobic modification on carbon dioxide absorption using porous alumina (Al2O3) hollow fiber membrane contactor," Journal of Membrane Science, vol. 518, pp. 79-87, 2016/11/15/ 2016, doi: https://doi.org/10.1016/j.memsci.2016.06.038.

[159]	F. Sandra et al., "Silicon carbide-based membranes with high soot particle filtration efficiency, durability and catalytic activity for CO/HC oxidation and soot combustion," Journal of Membrane Science, vol. 501, pp. 79-92, 2016/03/01/ 2016, doi: https://doi.org/10.1016/j.memsci.2015.12.015.

[160]	S. S. Manickam, J. Gelb, and J. R. McCutcheon, "Pore structure characterization of asymmetric membranes: Non-destructive characterization of porosity and tortuosity," Journal of Membrane Science, vol. 454, pp. 549-554, 2014/03/15/ 2014, doi: https://doi.org/10.1016/j.memsci.2013.11.044.

[161]	E. Akhondi, F. Zamani, J. W. Chew, W. B. Krantz, and A. G. Fane, "Improved design and protocol for evapoporometry determination of the pore-size distribution," Journal of Membrane Science, vol. 496, pp. 334-343, Dec 2015, doi: 10.1016/j.memsci.2015.09.013.

[162]	E. Akhondi, F. Zamani, A. W. K. Law, W. B. Krantz, A. G. Fane, and J. W. Chew, "Influence of backwashing on the pore size of hollow fiber ultrafiltration membranes," Journal of Membrane Science, vol. 521, pp. 33-42, Jan 2017, doi: 10.1016/j.memsci.2016.08.070.

[163]	F. Zamani, E. Akhondi, G. H. Koops, W. B. Krantz, A. G. Fane, and J. W. Chew, "Evapoporometry adaptation to determine the lumen-side pore-size distribution (PSD) of hollow fiber and tubular membranes," Journal of Membrane Science, vol. 526, pp. 1-8, Mar 2017, doi: 10.1016/j.memsci.2016.12.024.

[164]	F. Zamani, P. Jayaraman, E. Akhondi, W. B. Krantz, A. G. Fane, and J. W. Chew, "Extending the uppermost pore diameter measureable via Evapoporometry," Journal of Membrane Science, vol. 524, pp. 637-643, Feb 2017, doi: 10.1016/j.memsci.2016.11.082.

[165]	W. Krantz, A. Greenberg, E. Kujundzic, A. Yeo, and S. Hosseini, A Novel Technique for Characterizing Membrane Pore Size - Evapoporometry. 2010.

[166]	W. Krantz, A. Greenberg, E. Kujundzic, A. Yeo, and S. Hosseini, Evapoporometry – A Novel Method for Determining the Pore-Size Distribution of MF and UF Membranes. 2011.

[167]	E. Akhondi, F. Wicaksana, W. B. Krantz, and A. G. Fane, "Evapoporometry determination of pore-size distribution and pore fouling of hollow fiber membranes," Journal of Membrane Science, vol. 470, pp. 334-345, Nov 2014, doi: 10.1016/j.memsci.2014.07.042.

[168]	C. Eyraud, M. Betemps, J. F. Quinson, F. Chatelut, M. Brun, and B. Rasneur, "DETERMINATION OF THE PORE-SIZE DISTRIBUTION OF AN ULTRAFILTER BY - GAS-LIQUID PERMPOROMETRY MEASUREMENT - COMPARISON BETWEEN FLOW POROMETRY AND CONDENSATE EQUILIBRIUM POROMETRY," Bulletin De La Societe Chimique De France Partie I-Physicochimie Des Systemes Liquides Electrochimie Catalyse Genie Chimique, no. 9-10, pp. I237-I244, 1984. [Online]. Available: <Go to ISI>://WOS:A1984TX11700001.

[169]	A. Mey-Marom and M. G. Katz, "Measurement of active pore size distribution of microporous membranes - a new approach," Journal of Membrane Science, vol. 27, no. 2, pp. 119-130, 1986/06/01/ 1986, doi: https://doi.org/10.1016/S0376-7388(00)82049-9.

[170]	M. G. Katz and G. Baruch, "New insights into the structure of microporous membranes obtained using a new pore size evaluation method," Desalination, vol. 58, no. 3, pp. 199-211, 1986/01/01/ 1986, doi: https://doi.org/10.1016/0011-9164(86)87004-7.

[171]	F. P. Cuperus, D. Bargeman, and C. A. Smolders, "Characterization of anisotropic UF-membranes: top layer thickness and pore structure," Journal of Membrane Science, vol. 61, pp. 73-83, 1991/01/01/ 1991, doi: https://doi.org/10.1016/0376-7388(91)80007-S.

[172]	M. G. Liu, R. Ben Aim, and M. Mietton Peuchot, "Characterization of Inorganic Membranes by Permporometry Method: Importance of Non Equilibrium Phenomena," Key Engineering Materials, vol. 61-62, pp. 603-606, 1992, doi: 10.4028/www.scientific.net/KEM.61-62.603.

[173]	G. Z. Cao, J. Meijernik, H. W. Brinkman, and A. J. Burggraaf, "Permporometry study on the size distribution of active pores in porous ceramic membranes," Journal of Membrane Science, vol. 83, no. 2, pp. 221-235, 1993/08/26/ 1993, doi: https://doi.org/10.1016/0376-7388(93)85269-3.

[174]	F. P. Cuperus, D. Bargeman, and C. A. Smolders, "Permporometry: the determination of the size distribution of active pores in UF membranes," Journal of Membrane Science, vol. 71, no. 1, pp. 57-67, 1992/07/03/ 1992, doi: https://doi.org/10.1016/0376-7388(92)85006-5.

[175]	P. Huang, N. Xu, J. Shi, and Y. S. Lin, "Characterization of asymmetric ceramic membranes by modified permporometry," Journal of Membrane Science, vol. 116, no. 2, pp. 301-305, 1996/08/07/ 1996, doi: https://doi.org/10.1016/0376-7388(96)00097-X.

[176]	T. Tsuru, T. Hino, T. Yoshioka, and M. Asaeda, "Permporometry characterization of microporous ceramic membranes," Journal of Membrane Science, vol. 186, no. 2, pp. 257-265, 2001/05/30/ 2001, doi: https://doi.org/10.1016/S0376-7388(00)00692-X.

[177]	T. Tsuru, Y. Takata, H. Kondo, F. Hirano, T. Yoshioka, and M. Asaeda, "Characterization of sol–gel derived membranes and zeolite membranes by nanopermporometry," Separation and Purification Technology, vol. 32, no. 1, pp. 23-27, 2003/07/01/ 2003, doi: https://doi.org/10.1016/S1383-5866(03)00036-4.

[178]	N. Nishiyama, D.-H. Park, Y. Egashira, and K. Ueyama, "Pore size distributions of silylated mesoporous silica MCM-48 membranes," Separation and Purification Technology, vol. 32, no. 1, pp. 127-132, 2003/07/01/ 2003, doi: https://doi.org/10.1016/S1383-5866(03)00024-8.

[179]	S. R. Chowdhury, A. M. Peters, D. H. A. Blank, and J. E. t. Elshof, "Influence of porous substrate on mesopore structure and water permeability of surfactant templated mesoporous silica membranes," Journal of Membrane Science, vol. 279, no. 1, pp. 276-281, 2006/08/01/ 2006, doi: https://doi.org/10.1016/j.memsci.2005.12.015.

[180]	C. Wang, X. Liu, R. Cui, and B. Zhang, "In situ evaluation of defect size distribution for supported zeolite membranes," Journal of Membrane Science, vol. 330, no. 1, pp. 259-266, 2009/03/20/ 2009, doi: https://doi.org/10.1016/j.memsci.2008.12.061.

[181]	J. Hedlund, D. Korelskiy, L. Sandström, and J. Lindmark, "Permporometry analysis of zeolite membranes," Journal of Membrane Science, vol. 345, no. 1, pp. 276-287, 2009/12/01/ 2009, doi: https://doi.org/10.1016/j.memsci.2009.09.012.

[182]	S. Higgins, R. Kennard, N. Hill, J. DiCarlo, and W. J. DeSisto, "Preparation and characterization of non-ionic block co-polymer templated mesoporous silica membranes," Journal of Membrane Science, vol. 279, no. 1, pp. 669-674, 2006/08/01/ 2006, doi: https://doi.org/10.1016/j.memsci.2006.01.014.

[183]	S. R. Chowdhury, J. De Lamare, and V. Valtchev, "Synthesis and structural characterization of EMT-type membranes and their performance in nanofiltration experiments," Journal of Membrane Science, vol. 314, no. 1, pp. 200-205, 2008/04/30/ 2008, doi: https://doi.org/10.1016/j.memsci.2008.01.038.

[184]	G. C. C. Yang and C. J. Li, "Tubular TiO2/Al2O3 composite membranes: preparation, characterization, and performance in electrofiltration of oxide-CMP wastewater," Desalination, vol. 234, no. 1, pp. 354-361, 2008/12/31/ 2008, doi: https://doi.org/10.1016/j.desal.2007.09.104.

[185]	J. Kuhn, S. Sutanto, J. Gascon, J. Gross, and F. Kapteijn, "Performance and stability of multi-channel MFI zeolite membranes detemplated by calcination and ozonication in ethanol/water pervaporation," Journal of Membrane Science, vol. 339, no. 1, pp. 261-274, 2009/09/01/ 2009, doi: https://doi.org/10.1016/j.memsci.2009.05.006.

[186]	M. C. Schillo, I. S. Park, W. V. Chiu, and H. Verweij, "Rapid thermal processing of inorganic membranes," Journal of Membrane Science, vol. 362, no. 1, pp. 127-133, 2010/10/15/ 2010, doi: https://doi.org/10.1016/j.memsci.2010.06.030.

[187]	J. Lindmark et al., "Impregnation of zeolite membranes for enhanced selectivity," Journal of Membrane Science, vol. 365, no. 1, pp. 188-197, 2010/12/01/ 2010, doi: https://doi.org/10.1016/j.memsci.2010.09.006.

[188]	B. Hofs, J. Ogier, D. Vries, E. F. Beerendonk, and E. R. Cornelissen, "Comparison of ceramic and polymeric membrane permeability and fouling using surface water," Separation and Purification Technology, vol. 79, no. 3, pp. 365-374, 2011/06/24/ 2011, doi: https://doi.org/10.1016/j.seppur.2011.03.025.

[189]	S. Zeidler, P. Puhlfürß, U. Kätzel, and I. Voigt, "Preparation and characterization of new low MWCO ceramic nanofiltration membranes for organic solvents," Journal of Membrane Science, vol. 470, pp. 421-430, 2014/11/15/ 2014, doi: https://doi.org/10.1016/j.memsci.2014.07.051.

[190]	S. Karimi et al., "A simple method for blocking defects in zeolite membranes," Journal of Membrane Science, vol. 489, pp. 270-274, 2015/09/01/ 2015, doi: https://doi.org/10.1016/j.memsci.2015.04.038.

[191]	S. Blumenschein, A. Böcking, U. Kätzel, S. Postel, and M. Wessling, "Rejection modeling of ceramic membranes in organic solvent nanofiltration," Journal of Membrane Science, vol. 510, pp. 191-200, 2016/07/15/ 2016, doi: https://doi.org/10.1016/j.memsci.2016.02.042.

[192]	N. Chang, H. Tang, L. Bai, Y. Zhang, and G. Zeng, "Optimized rapid thermal processing for the template removal of SAPO-34 zeolite membranes," Journal of Membrane Science, vol. 552, pp. 13-21, 2018/04/15/ 2018, doi: https://doi.org/10.1016/j.memsci.2018.01.066.

[193]	A. Simon, H. Richter, B. Reif, M. Schuelein, D. Sanwald, and W. Schwieger, "Evaluation of a method for micro-defect sealing in ZSM-5 zeolite membranes by chemical vapor deposition of carbon," Separation and Purification Technology, vol. 219, pp. 180-185, 2019/07/15/ 2019, doi: https://doi.org/10.1016/j.seppur.2019.03.017.

[194]	K. S. W. Sing, "Physisorption of nitrogen by porous materials," Journal of Porous Materials, vol. 2, no. 1, pp. 5-8, 1995/07/01 1995, doi: 10.1007/BF00486564.

[195]	B. Nie, X. Liu, L. Yang, J. Meng, and X. Li, "Pore structure characterization of different rank coals using gas adsorption and scanning electron microscopy," Fuel, vol. 158, pp. 908-917, 2015/10/15/ 2015, doi: https://doi.org/10.1016/j.fuel.2015.06.050.

[196]	E. M. Cuerda-Correa, M. Díaz-Díez, A. Macías-García, and J. Gañán-Gómez, "Determination of the fractal dimension of activated carbons: Two alternative methods," Applied Surface Science - APPL SURF SCI, vol. 252, pp. 6102-6105, 06/01 2006, doi: 10.1016/j.apsusc.2005.11.012.

[197]	G.-J. Lee, S.-I. Pyun, and C. K. Rhee, "Characterisation of geometric and structural properties of pore surfaces of reactivated microporous carbons based upon image analysis and gas adsorption," Microporous and Mesoporous Materials - MICROPOROUS MESOPOROUS MAT, vol. 93, pp. 217-225, 07/01 2006, doi: 10.1016/j.micromeso.2006.02.025.

[198]	M. Schmitt, C. P. Fernandes, J. A. B. da Cunha Neto, F. G. Wolf, and V. S. S. dos Santos, "Characterization of pore systems in seal rocks using Nitrogen Gas Adsorption combined with Mercury Injection Capillary Pressure techniques," Marine and Petroleum Geology, vol. 39, no. 1, pp. 138-149, 2013/01/01/ 2013, doi: https://doi.org/10.1016/j.marpetgeo.2012.09.001.

[199]	P. Prádanos, M. L. Rodriguez, J. I. Calvo, A. Hernández, F. Tejerina, and J. A. de Saja, "Structural characterization of an UF membrane by gas adsorption-desorption and AFM measurements," Journal of Membrane Science, vol. 117, no. 1, pp. 291-302, 1996/08/21/ 1996, doi: https://doi.org/10.1016/0376-7388(96)00081-6.

[200]	C. R. Clarkson and R. Marc Bustin, "Variation in micropore capacity and size distribution with composition in bituminous coal of the Western Canadian Sedimentary Basin: Implications for coalbed methane potential," Fuel, vol. 75, no. 13, pp. 1483-1498, 1996/10/01/ 1996, doi: https://doi.org/10.1016/0016-2361(96)00142-1.

[201]	S. Brunauer, P. H. Emmett, and E. Teller, "Adsorption of Gases in Multimolecular Layers," Journal of the American Chemical Society, vol. 60, no. 2, pp. 309-319, 1938/02/01 1938, doi: 10.1021/ja01269a023.

[202]	E. P. Barrett, L. G. Joyner, and P. P. Halenda, "The Determination of Pore Volume and Area Distributions in Porous Substances. I. Computations from Nitrogen Isotherms," Journal of the American Chemical Society, vol. 73, no. 1, pp. 373-380, 1951/01/01 1951, doi: 10.1021/ja01145a126.

[203]	J. Seifert and G. Emig, "Mikrostrukturuntersuchungen an porösen Feststoffen durch Physisorptionsmessungen," Chemie Ingenieur Technik, vol. 59, no. 6, pp. 475-484, 1987/06/01 1987, doi: 10.1002/cite.330590606.

[204]	K. S. W. S. S. J. Gregg, Adsorption, Surface Area and Porosity (Academic Press, no. 10). London, UK: John Wiley & Sons, Ltd, 1982, pp. 957-957.

[205]	Y. Chen, L. Wei, M. Mastalerz, and A. Schimmelmann, "The effect of analytical particle size on gas adsorption porosimetry of shale," International Journal of Coal Geology, vol. 138, pp. 103-112, 2015/01/15/ 2015, doi: https://doi.org/10.1016/j.coal.2014.12.012.

[206]	K. Kaneko, "Determination of pore size and pore size distribution: 1. Adsorbents and catalysts," Journal of Membrane Science, vol. 96, no. 1, pp. 59-89, 1994/11/28/ 1994, doi: https://doi.org/10.1016/0376-7388(94)00126-X.

[207]	M. M. Dubinin, "The Potential Theory of Adsorption of Gases and Vapors for Adsorbents with Energetically Nonuniform Surfaces," Chemical Reviews, vol. 60, no. 2, pp. 235-241, 1960/04/01 1960, doi: 10.1021/cr60204a006.

[208]	M. M. Dubinin, "Fundamentals of the theory of adsorption in micropores of carbon adsorbents: Characteristics of their adsorption properties and microporous structures," Carbon, vol. 27, no. 3, pp. 457-467, 1989/01/01/ 1989, doi: https://doi.org/10.1016/0008-6223(89)90078-X.

[209]	M. M. Dubinin, "Fundamentals of the theory of adsorption in micropores of carbon adsorbents: characteristics of their adsorption properties and microporous structures," Pure and Applied Chemistry, vol. 61, no. 11, pp. 1841-1843, 1989, doi: 10.1351/pac198961111841.

[210]	A. Gil, A. Díaz, M. Montes, and D. R. Acosta, "Characterization of the microporosity of pillared clays by nitrogen adsorption — application of the Horvath-Kawazoe approach," Journal of Materials Science, journal article vol. 29, no. 18, pp. 4927-4932, September 01 1994, doi: 10.1007/bf00356545.

[211]	G. Horváth, Kamazoe, K., "METHOD FOR THE CALCULATION OF EFFECTIVE PORE SIZE DISTRIBUTION IN MOLECULAR SIEVE CARBON," Journal of Chemical Engineering of Japan, vol. 16, no. 6, pp. 470-475, 1983, doi: 10.1252/jcej.16.470.

[212]	B. C. Lippens and J. H. de Boer, "Studies on pore systems in catalysts: V. The t method," Journal of Catalysis, vol. 4, no. 3, pp. 319-323, 1965/06/01/ 1965, doi: https://doi.org/10.1016/0021-9517(65)90307-6.

[213]	W. D. Harkins and G. Jura, "Surfaces of Solids. XII. An Absolute Method for the Determination of the Area of a Finely Divided Crystalline Solid," Journal of the American Chemical Society, vol. 66, no. 8, pp. 1362-1366, 1944/08/01 1944, doi: 10.1021/ja01236a047.

[214]	L. Palacio, P. Prádanos, J. Calvo, and A. Hernández, "Porosity Determinations by a New Gas Penetration Volumetric Method," Acta Polithecnica Scandinavica. Chemical Technology Series, vol. 247, pp. 115-120, 01/01 1997.

[215]	L. Palacio, P. Prádanos, J. I. Calvo, and A. Hernández, "Porosity measurements by a gas penetration method and other techniques applied to membrane characterization," Thin Solid Films, vol. 348, no. 1, pp. 22-29, 1999/07/06/ 1999, doi: https://doi.org/10.1016/S0040-6090(99)00197-2.

[216]	J. I. Calvo et al., "Bulk and surface characterization of composite UF membranes Atomic force microscopy, gas adsorption-desorption and liquid displacement techniques," Journal of Membrane Science, vol. 128, no. 1, pp. 7-21, 1997/05/28/ 1997, doi: https://doi.org/10.1016/S0376-7388(96)00304-3.

[217]	Y. Fang, L. Bian, Q. Bi, Q. Li, and X. Wang, "Evaluation of the pore size distribution of a forward osmosis membrane in three different ways," Journal of Membrane Science, vol. 454, pp. 390-397, 2014/03/15/ 2014, doi: https://doi.org/10.1016/j.memsci.2013.12.046.

[218]	K. Hazazi et al., "Ultra-selective carbon molecular sieve membranes for natural gas separations based on a carbon-rich intrinsically microporous polyimide precursor," Journal of Membrane Science, vol. 585, pp. 1-9, 2019/09/01/ 2019, doi: https://doi.org/10.1016/j.memsci.2019.05.020.

[219]	A. Hernández, J. I. Calvo, P. Prádanos, and F. Tejerina, "Pore size distributions of track-etched membranes; comparison of surface and bulk porosities," Colloids and Surfaces A: Physicochemical and Engineering Aspects, vol. 138, no. 2, pp. 391-401, 1998/07/15/ 1998, doi: https://doi.org/10.1016/S0927-7757(96)03974-X.

[220]	Y. Sada et al., "Preparation and characterization of organic chelate ligand (OCL)-templated TiO2–ZrO2 nanofiltration membranes," Journal of Membrane Science, vol. 591, p. 117304, 2019/12/01/ 2019, doi: https://doi.org/10.1016/j.memsci.2019.117304.

[221]	S. Dong et al., "Freeze casting of novel porous silicate cement supports using tert-butyl alcohol-water binary crystals as template: Microstructure, strength and permeability," Journal of Membrane Science, vol. 541, pp. 143-152, 2017/11/01/ 2017, doi: https://doi.org/10.1016/j.memsci.2017.06.067.

[222]	P. Cini, S. R. Blaha, M. P. Harold, and K. Venkataraman, "Preparation and characterization of modified tubular ceramic membranes for use as catalyst supports," Journal of Membrane Science, vol. 55, no. 1, pp. 199-225, 1991/01/01/ 1991, doi: https://doi.org/10.1016/S0376-7388(00)82335-2.

[223]	D. A. Bailey, C. D. Jones, A. R. Barron, and M. R. Wiesner, "Characterization of alumoxane-derived ceramic membranes," Journal of Membrane Science, vol. 176, no. 1, pp. 1-9, 2000/08/15/ 2000, doi: https://doi.org/10.1016/S0376-7388(00)00341-0.

[224]	A. K. Basumatary, R. V. Kumar, A. K. Ghoshal, and G. Pugazhenthi, "Synthesis and characterization of MCM-41-ceramic composite membrane for the separation of chromic acid from aqueous solution," Journal of Membrane Science, vol. 475, pp. 521-532, 2015/02/01/ 2015, doi: https://doi.org/10.1016/j.memsci.2014.10.055.

[225]	A. Maghsodi, L. Adlnasab, M. Shabanian, and M. Javanbakht, "Optimization of effective parameters in the synthesis of nanopore anodic aluminum oxide membrane and arsenic removal by prepared magnetic iron oxide nanoparicles in anodic aluminum oxide membrane via ultrasonic-hydrothermal method," Ultrasonics Sonochemistry, vol. 48, pp. 441-452, 2018/11/01/ 2018, doi: https://doi.org/10.1016/j.ultsonch.2018.07.003.

[226]	B. Zhang, Y. Shi, Y. Wu, T. Wang, and J. Qiu, "Preparation and characterization of supported ordered nanoporous carbon membranes for gas separation," Journal of Applied Polymer Science, vol. 131, no. 4, 2014/02/15 2014, doi: 10.1002/app.39925.

[227]	Y. Zeng, L. Prasetyo, S. J. Tan, C. Fan, D. D. Do, and D. Nicholson, "On the hysteresis of adsorption and desorption of simple gases in open end and closed end pores," Chemical Engineering Science, vol. 158, pp. 462-479, 2017/02/02/ 2017, doi: https://doi.org/10.1016/j.ces.2016.10.048.

[228]	C. Jallut, J. Lenoir, C. Bardot, and C. Eyraud, "Thermoporometry.: Modelling and simulation of a mesoporous solid," Journal of Membrane Science, vol. 68, no. 3, pp. 271-282, 1992/04/15/ 1992, doi: https://doi.org/10.1016/0376-7388(92)85028-H.

[229]	L. Zeman, G. Tkacik, and P. Le Parlouer, "Characterization of porous sublayers in uf membranes by thermoporometry," Journal of Membrane Science, vol. 32, no. 2, pp. 329-337, 1987/07/01/ 1987, doi: https://doi.org/10.1016/S0376-7388(00)85015-2.

[230]	J. F. Quinson, N. Mameri, L. Guihard, and B. Bariou, "The study of the swelling of an ultrafiltration membrane under the influence of solvents by thermoporometry and measurement of permeability," Journal of Membrane Science, vol. 58, no. 2, pp. 191-200, 1991/05/15/ 1991, doi: https://doi.org/10.1016/S0376-7388(00)82455-2.

[231]	J. F. Quinson, N. Nameri, and B. Bariou, "Textural Characterization of Ultrafiltration Membranes by Thermoporometry and Liquid Flow Measurement," in Studies in Surface Science and Catalysis, vol. 62, F. Rodriguez-Reinoso, J. Rouquerol, K. S. W. Sing, and K. K. Unger Eds.: Elsevier, 1991, pp. 209-216.

[232]	A. P. Broek, H. A. Teunis, D. Bargeman, E. D. Sprengers, and C. A. Smolders, "Characterization of hollow fiber hemodialysis membranes: pore size distribution and performance," Journal of Membrane Science, vol. 73, no. 2, pp. 143-152, 1992/10/09/ 1992, doi: https://doi.org/10.1016/0376-7388(92)80124-3.

[233]	F. P. Cuperus, D. Bargeman, and C. A. Smolders, "Critical points in the analysis of membrane pore structures by thermoporometry," Journal of Membrane Science, vol. 66, no. 1, pp. 45-53, 1992/02/01/ 1992, doi: https://doi.org/10.1016/0376-7388(92)80090-7.

[234]	L. P. Terrazas-Bandala et al., "Influence of humidity, temperature, and the addition of activated carbon on the preparation of cellulose acetate membranes and their ability to remove arsenic from water," Journal of Applied Polymer Science, vol. 131, no. 8, 2014/04/15 2014, doi: 10.1002/app.40134.

[235]	S. H. Kim, S.-Y. Kwak, and T. Suzuki, "Positron Annihilation Spectroscopic Evidence to Demonstrate the Flux-Enhancement Mechanism in Morphology-Controlled Thin-Film-Composite (TFC) Membrane," Environmental Science & Technology, vol. 39, no. 6, pp. 1764-1770, 2005/03/01 2005, doi: 10.1021/es049453k.

[236]	Z. Chen, K. Ito, H. Yanagishita, N. Oshima, R. Suzuki, and Y. Kobayashi, "Correlation study between free-volume holes and molecular separations of composite membranes for reverse osmosis processes by means of variable-energy positron annihilation techniques," Journal of Physical Chemistry C, Article vol. 115, no. 37, pp. 18055-18060, 2011, doi: 10.1021/jp203888m.

[237]	T. Fujioka et al., "Rejection of small and uncharged chemicals of emerging concern by reverse osmosis membranes: The role of free volume space within the active skin layer," Separation and Purification Technology, Article vol. 116, pp. 426-432, 2013, doi: 10.1016/j.seppur.2013.06.015.

[238]	T. Fujioka et al., "Effect of heat treatment on fouling resistance and the rejection of small and neutral solutes by reverse osmosis membranes," Water Science and Technology: Water Supply, vol. 15, no. 3, pp. 510-516, 2015.

[239]	J. Lee, C. M. Doherty, A. J. Hill, and S. E. Kentish, "Water vapor sorption and free volume in the aromatic polyamide layer of reverse osmosis membranes," Journal of Membrane Science, Article vol. 425-426, pp. 217-226, 2013, doi: 10.1016/j.memsci.2012.08.054.

[240]	J. Fan et al., "Structure dependence of water vapor permeation in polymer nanocomposite membranes investigated by positron annihilation lifetime spectroscopy," Journal of Membrane Science, vol. 549, pp. 581-587, 2018/03/01/ 2018, doi: https://doi.org/10.1016/j.memsci.2017.12.046.

[241]	S.-J. Kim et al., "Characterization of pore size distribution (PSD) in cellulose triacetate (CTA) and polyamide (PA) thin active layers by positron annihilation lifetime spectroscopy (PALS) and fractional rejection (FR) method," Journal of Membrane Science, vol. 527, pp. 143-151, 2017/04/01/ 2017, doi: https://doi.org/10.1016/j.memsci.2016.12.064.

[242]	X. Ma, H. Wang, H. Wang, J. O. Brien-Abraham, and Y. S. Lin, "Pore structure characterization of supported polycrystalline zeolite membranes by positron annihilation spectroscopy," Journal of Membrane Science, vol. 477, pp. 41-48, 2015/03/01/ 2015, doi: https://doi.org/10.1016/j.memsci.2014.12.018.

[243]	A. Pifferi, "Radiation Sources: X-Ray Sources," in Reference Module in Materials Science and Materials Engineering: Elsevier, 2016.

[244]	C. Streli, P. Wobrauschek, and P. Kregsamer, "X-ray Fluorescence Spectroscopy, Applications," in Encyclopedia of Spectroscopy and Spectrometry, J. C. Lindon Ed. Oxford: Elsevier, 1999, pp. 2478-2487.

[245]	C. R. A. Catlow, "Synchrotron radiation techniques in materials and environmental science," Philosophical Transactions of the Royal Society A: Mathematical, Physical and Engineering Sciences, vol. 373, no. 2036, p. 20130162, 2015, doi: doi:10.1098/rsta.2013.0162.

[246]	M. Lassinantti Gualtieri et al., "Crack formation in α-alumina supported MFI zeolite membranes studied by in situ high temperature synchrotron powder diffraction," Journal of Membrane Science, vol. 290, no. 1, pp. 95-104, 2007/03/01/ 2007, doi: https://doi.org/10.1016/j.memsci.2006.12.018.

[247]	H.-K. Jeong, Z. Lai, M. Tsapatsis, and J. C. Hanson, "Strain of MFI crystals in membranes: An in situ synchrotron X-ray study," Microporous and Mesoporous Materials, vol. 84, no. 1, pp. 332-337, 2005/09/15/ 2005, doi: https://doi.org/10.1016/j.micromeso.2005.05.043.

[248]	R. Barranco-García, J. M. López-Majada, V. Lorenzo, J. M. Gómez-Elvira, E. Pérez, and M. L. Cerrada, "Confinement of iPP chains in the interior of SBA-15 mesostructure ascertained by gas transport properties in iPP-SBA-15 nanocomposites prepared by extrusion," Journal of Membrane Science, Article vol. 569, pp. 137-148, 2019, doi: 10.1016/j.memsci.2018.10.009.

[249]	H. Grigoriew, S. Bernstorff, A. Wolińska-Grabczyk, J. Domagała, and A. G. Chmielewski, "Depth-influenced structure through permeating polymer membrane using SAXS synchrotron method," Journal of Membrane Science, vol. 186, no. 1, pp. 1-8, 2001/05/15/ 2001, doi: https://doi.org/10.1016/S0376-7388(00)00652-9.

[250]	H. Grigoriew, A. Wolinska-Grabczyk, A. G. Chmielewski, H. Amenitsch, and S. Bernstorff, "SAXS study of the influence of ethanol on the microstructure of polyurethane-based membrane," Journal of Membrane Science, vol. 170, no. 2, pp. 275-279, 2000/05/31/ 2000, doi: https://doi.org/10.1016/S0376-7388(99)00364-6.

[251]	J.-C. Remigy and M. Meireles, "Assessment of pore geometry and 3-D architecture of filtration membranes by synchrotron radiation computed microtomography," Desalination, vol. 199, no. 1, pp. 501-503, 2006/11/20/ 2006, doi: https://doi.org/10.1016/j.desal.2006.03.193.

[252]	S.-H. Lee, W.-S. Chang, S.-M. Han, D.-H. Kim, and J.-K. Kim, "Synchrotron X-ray nanotomography and three-dimensional nanoscale imaging analysis of pore structure-function in nanoporous polymeric membranes," Journal of Membrane Science, vol. 535, pp. 28-34, 2017/08/01/ 2017, doi: https://doi.org/10.1016/j.memsci.2017.04.024.

[253] Joekar-Niasar, V., S. M. Hassanizadeh et al. (2012) Analysis of Fundamentals of Two-Phase Flow in Porous Media Using Dynamic Pore-Network Models: A Review, J. Crit. Rev. Environ. Sci. Technol. vol. 42, 1895-1976.

[254] R Lenormand, E Touboul, C Zarcone, Numerical models and experiments on immiscible displacements in porous media, Journal of fluid mechanics, 189, 165-187 (1988).

[255] R. Chandler, Capillary displacement and percolation in porous media, J. Fluid Mech. (1982), vol. 119, 248-267.

[256] D. Stau_er, A. Aharony, Introduction to percolation theory, Taylor & Francis, London, 1992.

FICPM –

[257] R. Peinador-Dàvila , O. Maalal , M.B Tanis-Kanbur, J. W. Chew, D. Lasseux, M. Prat Liquid-Liquid Porometry, Evapoporometry and Pore Network Simulations of Ultrafiltration membranes. “13th World Filtration Congress”, San Diego USA (April 2020).